
[image: image1.jpg]


[image: image47.png]


[image: image48.jpg]



This page has been left blank intentionally.

eriss HPGe detector calibration

J Pfitzner
Supervising Scientist Division

GPO Box 461, Darwin NT 0801

October 2010
Registry File SSD2002/0216
(Release status – unrestricted)
[image: image49.jpg]Australian Government

Department of Sustainability, Environment,
Water, Population and Communities
Supervising Sclentist





How to cite this report:
Pfitzner J 2010. eriss HPGe detector calibration. Internal Report 576, October, Supervising Scientist, Darwin. Unpublished paper.
Location of final PDF file in SSDX Sharepoint:
Supervising Scientist Division > PublicationWork > Publications and Productions > Internal Reports (IRs) > Nos 500 to 599 > IR576_HPGe detector calibration (Pfitzner)
Location of all key data files for this report in SSDX Sharepoint: 

Supervising Scientist Division > SSDX > Equipment and Techniques > HPGe Gamma Spec > Calibrations > HPGe Cal IR Files
Author of this report:

John Pfitzner – Environmental Research Institute of the Supervising Scientist, GPO Box 461, Darwin NT 0801, Australia. 
Additional contact: Andreas Bollhöfer, Program Leader, Environmental Radioactivity, Environmental Research Institute of the Supervising Scientist, GPO Box 461, Darwin NT 0801, Australia.
The Supervising Scientist is part of the Australian Government Department of Sustainability, Environment, Water, Population and Communities.
© Commonwealth of Australia 2010
Supervising Scientist

Department of Sustainability, Environment, Water, Population and Communities 
GPO Box 461, Darwin NT 0801 Australia

This work is copyright. Apart from any use as permitted under the Copyright Act 1968, no part may be reproduced by any process without prior written permission from the Supervising Scientist. Requests and inquiries concerning reproduction and rights should be addressed to Publications Inquiries, Supervising Scientist, GPO Box 461, Darwin NT 0801. 

e-mail: publications_ssd@environment.gov.au

Internet: www.environment.gov.au/ssd   (www.environment.gov.au/ssd/publications)

The views and opinions expressed in this report do not necessarily reflect those of the Commonwealth of Australia. While reasonable efforts have been made to ensure that the contents of this report are factually correct, some essential data rely on references cited and/or the data and/or information of other parties, and the Supervising Scientist and the Commonwealth of Australia do not accept responsibility for the accuracy, currency or completeness of the contents of this report, and shall not be liable for any loss or damage that may be occasioned directly or indirectly through the use of, or reliance on, the report. Readers should exercise their own skill and judgment with respect to their use of the material contained in this report. 

Printed and bound in Darwin NT by Supervising Scientist Division

Contents

ivExecutive summary


1Introduction


2The calibration and calculation procedures


2Calibration flow chart


3Calculations in Visual Gamma


5Flow chart of initial calculations in Excel spreadsheets on ‘P’ file data


7Calculations in the Excel spreadsheet


9238U calibration factor standards and 226Ra ratio standards


10Flow chart of radium (226Ra) ratio calculations


11Flow chart of calculation for 238U series and 235U calibration


15232Th calibration factor standards and ratio standards


16Flow chart of 232Th series calibration


18Flow chart for 232Th ratios at 63keV


19Flow chart for 40K and 137Cs calibration


21Entering calibration data into Visual Gamma


22Validation procedure


23Other Quality Assurance Procedures


24References


25Appendix 1  Table of geometry designation, size, containment material, sample type and sample weight range


26Appendix 2  Current Calibration Data


42Appendix 3  Plots of absolute efficiencies for Detectors G, N and O


50Appendix 4  Tables of absolute efficiency values for detectors G, N and O


56Appendix 5  

 HYPERLINK \l "_Toc276479903" 
Listing of Gamma Spectroscopy Standards Used for calibration




Executive summary
The Environmental Radioactivity Group (EnRad) of the Supervising Scientist Division (SSD) has undertaken a comprehensive calibration of their High Purity Germanium (HPGe) gamma spectroscopy system. 

Three detectors have been calibrated for five different geometries (sample types) holding sample masses of less than a gram up to 150 grams. This IR documents the calibration procedure, the methods of calculation and the results of the system calibration. 

eriss HPGe detector calibration

J Pfitzner

Introduction

The Environmental Radioactivity (EnRad) Group in the Supervising Scientist Division (SSD) operates several High Purity Germanium (HPGe) detectors for the determination of radionuclide activity concentration by gamma spectroscopy analysis. Samples are analysed for naturally occurring radionuclides of the uranium and thorium decay series and 40K, as well as for the nuclear bomb fallout isotope 137Cs. These analyses are done for internal research projects and on a consultancy basis for external clients. 

For quality assurance purposes, routine background and calibration checks are conducted to ensure that the instruments are operating within specifications, and to ensure that accurate results for these analyses are reported. Replacing hardware or making adjustments to instrument settings can cause significant changes to the calibration. Recent changes to the hardware triggered the need to do the calibration described in this document. All three detectors (detectors N, G and O) in current use were calibrated. 

Samples are prepared in one of several ‘geometries’. Each geometry is assigned an identifying letter. The ‘geometries’ are designed to ensure consistentcy in the shape and quantity of sample presented to the detector. The samples are either confined within a plastic container (geometries ‘Q’ and ‘R’) or cast in polyester resin (geometries ‘A’, ‘B’ and ‘V’). The calibration was done for all five geometries for each of the three detectors.

The range of geometries allows for optimisation of counting for a variety of sample sizes and types. Geometries ‘A’ and ‘V’ are ideal for organic material, such as dried or ashed leaf litter or mussel flesh. Geometries ‘B’, ‘Q’ and ‘R’ are for sediments and soils and similar inorganic materials. For a more complete description of these geometries, see Appendix 1 Table of Geometry Designations, Sizes and Sample Types. 

Specific radionuclide standards are prepared for each geometry. The standards are grouped in their naturally occurring 238U and 232Th series. Separate standards are prepared for 40K and 137Cs. To reduce statistical uncertainty, the activity of the standards are such that a one day count will ensure an uncertainty of less than 1% in counting statistics alone. This level of activity also ensures that the instrument background count rate is less than 0.5% of the standard activity count rate. 

Standards are prepared in batches of three, to allow measurement of systematic uncertainty. Standards are typically counted for one day, whereas backgrounds are counted for three days, usually over a weekend. 

The calibration and calculation procedures

Calibration flow chart


[image: image2]
Calculations in Visual Gamma

The program Visual Gamma developed in-house is used to operate the detector systems; starting and stopping counting, as well as performing the spectrum analyses. The Visual Gamma Operator’s Manual (Esparon & Pfitzner 2010) describes the calculation procedures for radionuclide analysis in Appendices B: Sources of uncertainty and propagation of uncertainty equations, C: Equations for the calculation of weighted average and D: Peak area and uncertainties by end point averaging. 

For calibration, the results of the calculation of standard and matrix and/or instrument background count rates, less continuum background, are required. At an intermediate stage of the analysis process, Visual Gamma produces a ‘P’ file which contains the desired count data, and count time in kiloseconds, for calibration calculations. Fig. 1 shows a typical ‘P’ file.

These ‘P’ files are read into an Excel spreadsheet, in which all relevant counting details are retained for Quality Assurance. Isotope line count data and associated uncertainty values are copied into the Excel sheet columns for manipulation. The copying process does not separate the ‘live’ count time of the standard or background from text, so this value , in ksec, must be manually entered in separate cells, identified by a label ‘ksec’ in an adjacent cell. 

The ‘P’ file count data for each isotope line are :

(1) the end point averaged counts for the peak region of interest (ROI) range of channels, with the continuum, ie the average counts of the left and right background ROI channels adjacent to the peak, subtracted.


[image: image3.wmf]÷

÷

ø

ö

ç

ç

è

æ

+

-

=

2

2

1

1

2

n

B

n

B

n

P

A

 

where A is the end point averaged counts, P is the total counts within the peak ROI of n channels, and B1 and B2 are the total counts in the adjacent background ROIs of n1 and n2 channels respectively.

(2) the 1 standard deviation uncertainty associated with the above calculations is calculated using:


[image: image4.wmf]ú

û

ù

ê

ë

é

÷

÷

ø

ö

ç

ç

è

æ

+

÷

ø

ö

ç

è

æ

+

=

2

2

2

2

1

1

2

2

n

B

n

B

n

P

A

s

 

Usually three standards of each radionuclide series in each geometry are counted on each detector. In some cases there is only one standard for a particular radionuclide in that geometry, eg 137Cs in A, B and V geometries.

The Excel worksheets for all the ‘P’ file results for one radionuclide series (U, Th, 40K or 137Cs) in one geometry (A, B, V, Q or R) for one detector (N, G or O) are in one Excel file. For example Det-G-Geom-Q-Th-Cal.xls contains the calibration data for detector ‘G’, geometry ‘Q’, 232Th calibration. This naming convention is maintained for all detectors, geometries and radionuclides.

There are two sets of ratios calculated, the ratio of 232Th in 234Th at the 63 keV energy line and the ratio of 226Ra in 235U at the 186keV energy line. The ‘P’ files of ratio calculations are incorporated into the respective Th and U series Excel spreadsheets.

Latest calibration data are stored on SSDX Sharepoint in the following location: SSDX Equipment and Techniques > HPGe Gamma Spec > Calibrations > (year).
Spectrum File: ZV034o.01

Description  : ZV034

Cal File     : Cal9o.100

Analyst      : John Pfitzner

MCB          : 7

Start Time   : 4/05/2010 1:13:42 PM

Live Time    : 77388

Mass         : 0.001996 kg

                                        Sample Counts

      

    Test Source Counts

Radio
   Energy
Peak Area

1 StdDev

Peak Area

1 StdDev

Ra-223
   270

217.50


78.52


733.50


40.80



Th-234
   63

-65.83


112.15


65,016.50

272.16



Pa-234
   1001

-31.44


35.58


3,126.50

56.42



Th-230
   68

261.86


99.44


3,180.50

105.58



Ra-226
   186

11,098.73

140.27


43,098.50

214.03



Pb-214
   295

35,460.16

200.23


747.00


37.81



Pb-214
   352

56,213.59

248.94


906.50


37.65



Bi-214
   609

33,731.99

190.56


482.00


26.79



Pb-210
   46

9,463.24

132.25


37,210.00

210.01



Ac-228
   338

68.00


64.59


297.50


26.86



Ac-228
   911

-51.17


44.85


51.50


10.52



Ac-228
   969

-31.38


38.97


23.00


8.34



Ra-224
   241

16,375.51

147.57


770.50


42.23



Pb-212
   239

102.31


75.74


703.00


40.74



Bi-212
   40

104.35


79.36


2,281.00

87.77



Tl-208
   583

72.18


48.19


175.00


20.22



K-40
   1461

5.18


27.41


5.00


3.94



Cs-137*   662

-25.18


45.12


143.50


19.06



Be-7
   478

-68.07


51.16


174.50


23.39


Radio

Energy

PEAK

   Left Back
 
  Right Back

Ra-223

270

1428
1449
1418
1427
1493
1502


Th-234

63

347

360

341

346

361

367


Pa-234

1001
5328
5347
5319
5327
5372
5380


Th-230

68

371

382

364

370

383

389


Ra-226

186

1000
1016
990

999

1017
1025


Pb-214

295

1578
1596
1568
1577
1630
1638


Pb-214

352

1878
1900
1867
1877
1901
1911


Bi-214

609

3243
3267
3232
3242
3268
3279


Pb-210

46

261

273

254

260

274

279


Ac-228

338

1807
1824
1799
1806
1837
1845


Ac-228

911

4846
4871
4801
4813
4872
4883


Ac-228

969

5154
5178
5119
5131
5179
5191


Ra-224

241

1297
1313
1260
1267
1314
1322


Pb-212

239

1277
1293
1257
1264
1314
1321


Bi-212

40

224

234

217

223

235

240


Tl-208

583

3105
3127
3080
3090
3128
3138


K-40

1461
7764
7784
7753
7763
7785
7795


Cs-137*
662

3521
3544
3510
3520
3563
3574


Be-7

478

2547
2567
2535
2546
2616
2628

Figure 1  A typical ‘P’ file, for radium ratio standard SV034 counted on detector ‘O’
Flow chart of initial calculations in Excel spreadsheets on ‘P’ file data

The instructions flowsheet below assumes that the operator uses the template files that have been previously prepared for the calibration process. These are located on the SSDX Sharepoint path: SSDX Equipment and Techniques > HPGe Gamma Spec > Calibrations. SSDX Sharepoint is the Division’s electronic document management and retrieval system (EDRMS). A typical Excel worksheet for calculation of the calibration factor, Cts/Ksec/Bq for a 137Cs standard in ‘A’ geometry on detector ‘N’ is shown in Figure 2.


[image: image5]
	Spectrum File: ZA034n.01
	
	
	
	
	
	
	
	
	
	
	
	

	Description:ZA034
	
	
	
	
	
	
	
	
	
	
	
	
	

	Cal File     : Cal1n.103
	
	
	137Cs
	
	
	
	
	
	
	
	
	

	Analyst      : John Pfitzner
	
	
	43.71
	Tot Bq
	Mar2010
	
	
	
	
	
	
	

	MCB          : 1
	
	
	
	
	
	
	
	
	
	
	
	
	

	Start Time: 20/04/2010 2:08:43 PM
	
	
	
	
	
	
	
	
	
	
	

	Live Time: 79513
	
	
	ksec
	79.513
	ksec
	
	
	
	
	
	
	
	

	Mass: 13.6365 kg
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	81.27
	Bq 137Cs
	Mar1983
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	Sample Counts
	
	
	
	
	
	
	Cts/Ksec
	Cts/ksec
	1sd
	1sduncert
	Cts/Ksec/Bq
	1sd

	Radio
	
	Energy
	
	Peak Area
	1StdDev
	
	Peak
	less bgd
	uncert
	less bgd
	Peak
	
	uncert

	Ra-223
	
	270
	
	74.22
	
	65.81
	
	0.9334
	0.9871
	0.8277
	0.8368
	
	
	

	Th-234
	
	63
	
	129.06
	
	53.64
	
	1.6231
	1.0335
	0.6746
	0.6834
	
	
	

	Pa-234
	
	1001
	
	-1.64
	
	13.02
	
	-0.0206
	0.0704
	0.1637
	0.1719
	
	
	

	Th-230
	
	68
	
	-9
	
	51.64
	
	-0.1132
	-0.6902
	0.6495
	0.6587
	
	
	

	Ra-226
	
	186
	
	186.93
	
	82.69
	
	2.3509
	1.6793
	1.0400
	1.0518
	
	
	

	Pb-214
	
	295
	
	39.55
	
	61.36
	
	0.4974
	-0.0961
	0.7717
	0.7803
	
	
	

	Pb-214
	
	352
	
	46.39
	
	64.19
	
	0.5834
	-0.6407
	0.8073
	0.8146
	
	
	

	Bi-214
	
	609
	
	48.92
	
	31.06
	
	0.6152
	-0.3251
	0.3906
	0.4002
	
	
	

	Pb-210
	
	46
	
	-82
	
	59.73
	
	-1.0313
	-2.0136
	0.7512
	0.7596
	
	
	

	Ac-228
	
	338
	
	-13.1
	
	59.14
	
	-0.1648
	-0.3357
	0.7438
	0.7505
	
	
	

	Ac-228
	
	911
	
	-3.29
	
	14.67
	
	-0.0414
	-0.1680
	0.1845
	0.1934
	
	
	

	Ac-228
	
	969
	
	21.45
	
	14.84
	
	0.2698
	0.1486
	0.1866
	0.1952
	
	
	

	Ra-224
	
	241
	
	174.25
	
	64.18
	
	2.1915
	1.8962
	0.8072
	0.8167
	
	
	

	Pb-212
	
	239
	
	279.25
	
	64.99
	
	3.5120
	2.7238
	0.8174
	0.8272
	
	
	

	Bi-212
	
	40
	
	-618.56
	
	63.65
	
	-7.7794
	-7.7587
	0.8005
	0.8082
	
	
	

	Tl-208
	
	583
	
	35
	
	30.32
	
	0.4402
	0.0915
	0.3813
	0.3887
	
	
	

	K-40
	
	1461
	
	59.58
	
	11.4
	
	0.7493
	0.1953
	0.1434
	0.1509
	
	
	

	Cs-137*
	
	662
	
	75,676.18
	
	276.89
	
	951.7460
	951.6528
	3.4823
	3.4830
	21.7717
	137Cs
	0.0797

	Be-7
	
	478
	
	175.61
	
	60.28
	
	2.2086
	2.1342
	0.7581
	0.7617
	
	
	


6
Figure 2  Worksheet for calculation of the calibration factor, Cts/Ksec/Bq for a 137Cs standard in ‘A’ geometry on detector ‘N’
Calculations in the Excel spreadsheet

The calculations described below use the ‘net count rate’ of the standard. The net rate is the difference between the counts per kilosecond of the standard from the ‘P’ file (from which the continuum background has already been subtracted) and the background, also in units of counts per kilosecond. The background may be a count of either a matrix  material (eg Flying Fox sand or dried beef) or the instrument background. Matrix materials must be prepared in the same geometry and mass as the standard they represent. The 1 standard deviation (1 s.d.) uncertainty associated with this net count rate is calculated using the formula from the Visual Gamma Operator’s Manual (Esparon & Pfitzner 2010) Appendix B: Sources of uncertainty and propagation of uncertainty equations.

[image: image6.wmf]2

2

2

2

2

1

2

1

x

x

y

s

a

s

a

s

+

=


where 
[image: image7.wmf]2

2

1

1

x

a

x

a

y

-

=

 (x1 is uncertainty in a1 and x2 in a2 and sy is 1 s.d. uncertainty in y).

The total activity of the radionuclide in the standard is entered into a clearly labelled cell near the top of the worksheet. Particular attention must be paid to short lived isotopes, such as 137Cs, where corrections for decay must be applied. Details of all eriss standards, including total activity, is maintained in the Excel spreadsheet ‘All eriss gamma standards.xls’, located at SSDX>Equipment and Techniques > HPGe Gamma Spec> Calibrations>Standards Listing.

If the standard is to be used only for ratios, the activity value is not used. In this case the next calculation is to select the specific cells for the ratio calculations. These are:
a) For the ratio of 232Th in 234Th at the 63 keV energy line the ratios used are 

63keV:338keV

63keV:911keV and 

63keV:969keV.

b) For the ratio of 226Ra in 235U at the 186keV energy line the ratios used are

186keV:295keV

186keV:352keV and 

186keV:609keV.
	238U series
	

	Isotope
	Line Energy (keV)

	235U
	185.720

	234Th
	63.300

	234Pa
	1001*

	230Th
	68*

	226Ra
	186.211

	214Pb
	295.224

	214Pb
	351.932

	214Bi
	609.312

	210Pb
	46.539

	214Pb
	241.997


The isotopes and energy lines for calibration of radionuclides in the U decay series are (see right):
(All line energies, except those marked with *, are obtained from the web site of the Nucleide Organisation, which maintains the latest updates of radionuclide data : http://www.nucleide.org/DDEP_WG/DDEPdata.htm)
The isotopes and energy lines for calibration of radionuclides in the Th-series are:

	232Th series
	

	Isotope
	Line Energy (keV)

	228Ac
	338.320

	228Ac
	911.196

	228Ac
	968.960

	224Ra
	240.986

	212Pb
	238.623

	212Bi
	39.858

	208Tl
	583.187


The isotopes and energy lines for calibration of radionuclides of the other routinely measured radionuclides are:

	Nuclides
	

	Isotope
	Line Energy (keV)

	40K
	1460.822

	137Cs (mBa)
	661.659


For each radionuclide line, the net count rate is divided by the total activty of that isotope in the standard. The result is counts per kilo second per Bq, the activity conversion factor. Uncertainties are propagated according to the formula found in the Visual Gamma Operator’s Manual (Esparon & Pfitzner 2010) Appendix B. ‘Sources of uncertainty and propagation of uncertainty equations’,


[image: image8.wmf]2

2

2

1

2

1

÷

÷

ø

ö

ç

ç

è

æ

+

÷

÷

ø

ö

ç

ç

è

æ

=

x

s

x

s

y

s

x

x

y


where 
[image: image9.wmf]2

1

x

x

y

=

.

A summary worksheet contains the weighted average of the activity conversion factor for several standards (b, c, and d) of the same type in the same geometry on the same detector. The formula for calculating the weighted average, Wm, is:


[image: image10.wmf]2

2

2

2

2

2

1

1

1

d

c

b

d

d

c

c

b

b

W

m

d

d

d

d

d

d

+

+

+

+

=


where the values and their associated uncertainties are b ± δb, c ±  δc and d ± δd, 

and the propagated uncertainty in the weighted mean is


[image: image11.wmf]2

2

2

1

1

1

1

d

c

b

UncertW

m

d

d

d

+

+

=

 
The weighted averages and uncertainties for the isotope lines are tabulated for each radionuclide line in the series. The summary page also containes the weighted averages of the ratios for that series.

The systematic uncertainty is now added to the propagated uncertainty using the formula 
[image: image12.wmf]2

2

2

2

2

1

2

1

x

x

y

s

a

s

a

s

+

=

from above. 

These systematic uncertainties are (Marten 1992):

Uranium (238U) series

1.2%

Thorium (232Th) series

1.5%

Others (137Cs & 40K)

2%

238U calibration factor standards and 226Ra ratio standards

eriss uses Canmet BL-5 as the standard reference material for 238U series calibration. BL-5 is a uranium ore in secular equilibrium with its progeny and the activity concentration of BL-5 is accurately and precisely known. 

The source material for 226Ra can be one of several 226Ra solutions which must not contain any 238U uranium.

Flow chart of radium (226Ra) ratio calculations

The 226Ra ratio must be completed before calculating the 238U calibration series as this ratio is used in the 235U calibration at 186keV. 


[image: image13]

Flow chart of calculation for 238U series and 235U calibration



	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	Cts/Ksec/Bq
	
	
	
	
	
	

	
	Peak
	
	uncert
	
	
	
	
	

	
	0.739585
	
	0.010612
	
	
	
	
	

	
	5.459564
	
	0.021811
	
	
	
	
	

	
	0.152224
	
	0.004682
	
	From radium ratios
	
	

	
	0.659348
	
	0.01675
	
	226Ra Contribution to 186kev
	

	
	6.401413
	
	0.020378
	
	Cts/ksec
	
	Uncert
	

	
	11.87108
	
	0.024049
	
	1204.9
	
	5.644739
	

	
	19.21909
	
	0.02996
	
	1227.086
	
	5.314179
	

	
	11.74032
	
	0.0232
	
	1228.531
	
	5.715505
	

	
	5.481455
	
	0.020314
	
	Wtd Ave.
	
	
	

	
	0.029694
	
	0.007917
	
	1220.392
	
	3.204102
	

	
	0.000369
	
	0.005096
	
	
	
	
	

	
	0.005775
	
	0.004384
	
	Net 235U at 186keV
	
	

	
	5.614979
	
	0.018243
	
	Cts/ksec
	
	uncert
	

	
	0.034385
	
	0.009511
	
	486.2248
	
	6.307214
	

	
	0.049522
	
	0.012998
	
	
	
	
	

	
	-0.00365
	
	0.004968
	
	235U at 186keV
	
	

	
	-0.00078
	
	0.003335
	
	Cts/ksec/Bq
	uncert
	

	
	0.001232
	
	0.004804
	
	1.823799
	
	0.023658
	

	
	0.006448
	
	0.006696
	
	
	
	
	

	
	
	
	
	
	226Ra at 186keV
	
	

	
	
	
	
	
	Cts/ksec/Bq
	uncert
	

	
	
	
	
	
	4.577614
	
	0.012018
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	


Figure 3  Excel worksheet showing the extra calculations required for 235U calibration factor at 186keV

	Det - G
	
	Geom - Q
	
	
	
	

	
	
	
	
	
	
	

	
	Radium Ratios:
	Combined
	Counting
	
	

	
	
	wtd ave
	uncert 1sd
	uncert 1sd
	
	

	
	186:295
	0.38072
	0.00161
	0.00161
	
	

	
	186:352
	0.23949
	0.00097
	0.00097
	
	

	
	186:609
	0.39251
	0.00165
	0.00165
	
	

	
	
	
	
	
	
	

	
	
	Using:
	ZQ056g.01
	ZQ055g.01
	ZQ019g.01
	

	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	Uranium Efficiencies:
	wtd ave
	Combined
	system
	counting

	
	
	
	Cts/ksec/Bq
	uncert 1sd
	uncert 1.2%
	uncert 1sd

	
	U-235
	186
	41.57010
	0.60936
	0.49884
	0.34997

	
	Th-234
	63
	5.41149
	0.06668
	0.06494
	0.01515

	
	Pa-234
	1001
	0.15271
	0.00373
	0.00183
	0.00324

	
	Th-230
	68
	0.62459
	0.01377
	0.00750
	0.01155

	
	Ra-226
	186
	4.43826
	0.05373
	0.05326
	0.00706

	
	Pb-214
	242
	5.45669
	0.06668
	0.06548
	0.01260

	
	Pb-214
	295
	11.51723
	0.13920
	0.13821
	0.01662

	
	Pb-214
	352
	18.68124
	0.22513
	0.22417
	0.02074

	
	Bi-214
	609
	11.38448
	0.13755
	0.13661
	0.01600

	
	Pb-210
	46
	5.43093
	0.06669
	0.06517
	0.01414

	
	
	
	
	
	
	

	
	
	Using Stds:
	ZQ049g.01
	ZQ001g.06
	ZQ003g.01
	ZQ050g.01

	
	
	
	
	
	
	

	
	Updated:
	235U @186 calib & uncert on 9 July 2010.
	

	
	
	
	
	
	
	

	
	on:
	
	
	
	
	

	
	
	
	
	
	
	

	
	Notes:
	
	
	
	
	

	
	
	
	
	
	
	

	
	Systematic uncertainties (Marten 1992)
	
	

	
	U: %
	1.2
	
	
	
	

	
	Th: %
	2
	
	
	
	

	
	all others %
	1.5
	
	
	
	

	
	
	
	
	
	
	


Figure 4  An Excel summary worksheet for U series isotopes on detector G, geometry Q. The highlighted cells are final values copied to the calibration page in Visual Gamma.

232Th calibration factor standards and ratio standards

The standards for 232Th ratio at 63keV must contain no uranium and the 232Th series progeny must be in isotopic equilibrium for the calibration factors. The only source material currently available at eriss which contains no U is an ‘Amersham 1906 thorium nitrate’ solution. Prior to 2005, the activity concentration of this solution was considered accurately known and the solution was used for 232Th series calibration factor standards as well as the ratio. Since 2005, due to a lack of confidence in the activity concentration, this solution has only been used to make 232Th ratio standards. The current source material for 232Th series calibration factor standards is the IAEA reference material RGTH-1 which contains a small amount of uranium and is therefore not suitable for the 232Th ratios at 63keV.

Flow chart of 232Th series calibration
This flow chart follows the initial calculations in the ‘P’ spreadsheet, see ‘Flow Chart of Initial Calculations in Excel spreadsheets on ‘P’ file data’ above.

	Det - G
	
	Geom - Q
	
	
	
	

	
	
	
	
	
	
	

	Thorium calibration
	
	
	
	
	

	
	
	
	Wtd ave
	Combined
	counting
	system

	
	Isotope
	line keV
	cts/ksec/Bq
	uncert 1sd
	uncert 1sd
	err 2%

	
	Ac-228
	338
	6.30702
	0.12691
	0.01399
	0.12614

	
	Ac-228
	911
	5.32166
	0.10709
	0.01182
	0.10643

	
	Ac-228
	969
	3.09670
	0.06259
	0.00907
	0.06193

	
	Ra-224
	241
	3.26282
	0.06627
	0.01154
	0.06526

	
	Pb-212
	239
	34.71357
	0.69493
	0.03029
	0.69427

	
	Bi-212
	40
	1.22936
	0.02659
	0.01011
	0.02459

	
	Tl-208
	583
	7.99019
	0.16050
	0.01497
	0.15980

	
	
	
	
	
	
	

	
	
	Using:
	ZQ008g.01&2
	ZQ005g.01&2
	ZQ054g.01
	ZQ065g.01

	
	
	
	
	
	
	

	
	Ratios
	
	Combined
	counting
	
	

	
	
	Wtd ave
	uncert 1sd
	uncert 1sd
	
	

	
	63:338
	0.05677
	0.00192
	0.00192
	
	

	
	63:911
	0.06732
	0.00228
	0.00228
	
	

	
	63:969
	0.11558
	0.00392
	0.00392
	
	

	
	
	
	
	
	
	

	
	
	
	Using:
	ZQ008g.01&2
	ZQ005g.01&2
	

	
	
	
	
	
	
	

	
	Updated:
	
	
	
	
	

	
	
	
	
	
	
	

	
	on:
	
	
	
	
	

	
	
	
	
	
	
	

	
	Notes:
	
	
	
	
	

	
	
	
	
	
	
	

	
	Systematic uncertainties (Marten 1992)
	
	
	

	
	U: %
	1.2
	
	
	
	

	
	Th: %
	2
	
	
	
	

	
	all others %
	1.5
	
	
	
	


Figure 5  An Excel summary worksheet for 232Th series isotopes on detector G, geometry Q. The highlighted cells are final values copied to the calibration page in Visual Gamma.
Flow chart for 232Th ratios at 63keV

This flow chart follows the initial calculations in the ‘P’ spreadsheet, see ‘Flow Chart of Initial Calculations in Excel spreadsheets on ‘P’ file data’ above.

[image: image14]
Flow chart for 40K and 137Cs calibration

Typically separate standards are made for 40K and 137Cs. There are several choices of source material for 40K as it is naturally occurring in potassium at the isotopic concentration of 0.001167%. Anatytical grade potassium salts, such as potassium sulphate and potassium nitrate are suitable sources. 137Cs, on the other hand, is a product of the nuclear testing during the period 1950 – 1990 and has been deposited in varying concentrations all over the earth’s terrestrial and aquatic surface. For 137Cs a source solution with a certified  activity concentration must be obtained from a reputable supplier, such as United Nuclear, GE Lifesciences or  Oak Ridge National Lab. Nuclear Science & Technology Division.

[image: image15]
	Det - G
	
	Geom - Q
	
	
	

	
	
	
	
	
	
	

	40K & 137Cs calibration
	
	
	
	

	
	
	
	wtd ave
	combined
	counting
	System

	
	Isotope
	line keV
	cts/ksec/Bq
	uncert 1sd
	uncert 1sd
	err 1.5%

	
	K-40
	1461
	1.38204
	0.03483
	0.02799
	0.0207306

	
	Cs-137
	662
	17.88292
	0.28395
	0.09314
	0.2682438

	
	
	
	
	
	
	

	
	
	
	Using :
	ZQ058g.02
	ZQ057g.01
	ZQ052g.01

	
	
	
	
	
	
	

	
	Entered:
	
	
	
	
	

	
	
	
	
	
	
	

	
	on:
	
	
	
	
	

	
	
	
	
	
	
	

	
	Notes:
	included systematic uncertainties from Marten 1992
	

	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	Systematic uncertainties (Marten 1992)
	
	

	
	U: %
	1.2
	
	
	
	

	
	Th: %
	2
	
	
	
	

	
	all others %
	1.5
	
	
	
	

	
	
	
	
	
	
	


Figure 6  Excel spreadsheet of Detector G, Geometry Q, summary sheet for 40K and 137Cs calibration
Entering calibration data into Visual Gamma
The following describes the step-by-step process of entering the calibration data into Visual Gamma. Copies of these calibration data pages are in Appendix 1. Current Calibration Data.


[image: image16]
Validation procedure

The validation procedure is an assurance check that the appropriate spectrum files have been used and calculations and data entry have been done correctly.



Other quality assurance procedures
The International Atomic Energy Agency conducts intercomparison excercises where laboratories are invited to participate in the radionuclide determination of specially prepared samples of environmental water, sediments or biota, with the aim of assessing the ability of each laboratory to accurately measure specific radionuclide activity concentrations. eriss have participated in several of these excercises during the past 20 years and have always achieved excellent accuracy and precision with their reported results. Following these intercomparison excercises, the sample material, with certified radionuclide concentrations, is made available for purchase as reference materials. Several of these reference materials are routinely analysed on the HPGe detectors to provide independent quality assurance of the eriss gamma spectroscopy calibrations and analytical procedures. The result of several years of counting these reference materials will be the subject of a separate Internal Report.

References

Esparon A & Pfitzner J (in press) Visual Gamma – Gamma Analysis Manual. Internal Report 539, Supervising Scientist, Darwin. 
Canberra Industries Product Catalogue, Edition twelve, p25.

Marten R 1992. Procedures for routine analysis of naturally occurring radionuclides in environmental samples by gamma-ray spectrometry with HPGe detectors. Internal report 76, Supervising Scientist for the Alligator Rivers Region, Canberra. Unpublished paper.

Appendix 1  Table of geometry designation, size, containment material, sample type and sample weight range
	Geometry Designation
	Size
height (mm) x radius (mm)
	Containment material
	Sample type
	Sample weight range
(g)

	A
	10 x 65
	polyester resin
	organic
	0 – 30

	B
	10 x 65
	polyester resin
	mineral
	10 – 30

	Q
	4 x 45
	acrylic
	mineral
	5 – 7

	R
	20 x 75
	acrylic
	mineral
	120 – 180

	V
	5 x 40
	polyester resin
	organic
	0 – 3


Appendix 2  Current calibration data

Detector G, Geometry A
27
[image: image17.jpg]B EfficiencyBackGround

Detector iG I Geometry |a a3 Current Setting 1lg.102
Etficiency Background ROI's o PTZIEETIIE
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)

e 270565 | [006552 [037e175 | (016174082 Relenaz6
Thaas 63 [sesesz | 00877 [a15  |[oraosos)  [33 | [s40 | | [m [ [ Hol ol
Thoas (825
Pazss ooy [oizes | oot [oossars | [oosssesze]  [ssos|[s;iz_| [s;13|[ssm | [s3|[sa
Tham (67 [nazs0e | (003425 [a022s | [oommossos]  [e [z | (36 (a4 | [a5s|[3m0
Re225 [125) [ze514 | [007171 [oweizs | jpimaoz  (se0 | [sem s |[woi | [100s | [1o12
Po2ts (202) [iaen Jormss ] [0 8|  rewmezs
Po2ne 295 [sains | o 55z | [ararem s [
Pozie @2 [1ssesr |[osres | [1aizs | [orrwsae e
6214 (e09) [ [0 1513278
Po210 45) 005775 (0756205 ]
Ac228 (3381 019115 (03925 | [orermiszz 170 | (1795
hozz (811) {4155 | (01655 [navizs | [ooszooo;n,  [a7es | [s800
et [239951 | 00360 [m1g7s | jporomssz  (me | [s11s
e (o8t | [o05m13 o o (245 ] 175
har ) [z7osmsa | [losoas | (2067 | [nssmeazz  [zes | [12%6
e e [om0s | [ooss0z [ooieiz | jposesizz (2| [zi6 ] ]
P [sas | [ozraz [oesesrs | [oizsoszz  [sovd | |30es 3| x|
ko (s [115 | [ooms [o4zs625 | [oosrzeres] 77w | [zraz | [z7a3_|[r7er | |[7ee |[77e |
e Bz [o00125 | [oo77aroes|  [aam0 | (309 | (30 || | [ass0 | (360
87 el (7055 | (00345 [o070125 | [oo720mo67| 516 | [2529 | [25%0 |[z5%6 | [a616 | (2620
Mess (535) (17086 | [o015 [oosters | [ooseororzl [ o | C—— 0
oz ) (006 | [000N1 | Razze (ee2ss) (0302096 | [000ess et
hoz2s (sa/a1n) (00531 |[0003%9 | mazos (iesse)  |016ssss | [0oozoss B e
Aoz e [ovsaws |[o00e17 | Razze (es/s03)  [0zserss | [noosses
Commarts
e Callo 100 odum atos erired WP 18 March 2010 Cal - Ok
Ca0 101 - updsted 13 Abn P} Gl DK

313102 Updatod 1375, 24 May 201DIJP), 40K 1o chock. Calb - OK.





Detector G, Geometry B
28
[image: image18.jpg]B EfficiencyBackGround

Detector iG 3 Geometry [ - Current Setting | 103
Efficiency Background ROI's EL
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) [ r—

e 200255 | [a0ss03 [037e175 | (016174082 Relenaz6

Thaas 63 [a7arsr | [o18161 [a15  |[oraosos)  [33 | [s40 | | [m [ [ Hol ol
Thoas (825

Fazss ooy (012 | (o007 [oossars | [oosssesze]  [ssos|[s;iz_| [s;13|[ssm | [s3|[sa

Tham (67 [naazes | (002575 [a022s | [oommossos]  [e [z | (36 (a4 | [a5s|[3m0

Re225 [125) [zeni2e | [o12102 [oweizs | jpimaoz  (se0 | [sem s |[woi | [100s | [1o12

Po2ts (202) [isest7 Jforzes ] [0 8 |  rewmezss

Po2ne 295 [szm0i | [oa0a 55z | [ararem s [

Poots s (s owen | iz | s —=

6214 (e09) [s5ase | | [0 1513278

e G - G .

Ac228 (3381 470265 | 013000 (03925 | [orermiszz 170 | (1795

hozz (811) (12833 | (016701 [navizs | [ooszooo;n,  [a7es | [s800

et [2425 | [00sees [m1g7s | jporomssz  (me | [s11s

e (391 | [o.09575 o o (245 ] 175

i [eszs0s | [1o7ere | (206w | joamamimz  [1zas |1

e e [o7reie | [oozen [ooieiz | jposesizz (2| [zi6 ] ]

P [emmz | [azezn [oesesrs | [oizsoszz  [sovd | |30es 3| x|

kio (s (112685 | [oomes [o4zs625 | [oosrzeres] 77w | [zraz | [z7a3_|[r7er | |[7ee |[77e |

e [1.92318 | 058557 [o00125 | [oo77aroes|  [aam0 | (309 | (30 || | [ass0 | (360

87 el {37055 | (00345 [o070125 | [oo720mo67| 516 | [2529 | [25%0 |[z5%6 | [a616 | (2620

Mess (535) (17086 | [o015 [oosters | [ooseororzl [ o | C—— 0

ac2z8 e [0 | (000541 | Rezz6 (186295 (03026 | [00ti6n et
hozzs (ga/ain) (005203 | (000814 | mazos (iess (0168 | [noone B e
o2 s [ovesse | [0mios1 | maze (w03 [0zaes | [ooiie
Commerts

Updated by AE to nclude Fa ratos. JP updated i atos (36228 & 5226] Uranum and Thorum series, 27 Aug 2008.
320102 Updsiod 15 Apil UF) 137C5 and 40K to chask. Cal - OK
€323 103  Updstod 40K & 1575 14 Apil UF). Caib - DR





Detector G, Geometry Q
29
[image: image19.jpg]B EfficiencyBackGround

Detector iG ~| Geometry g ~  Current Setting 105
Efficiency Background ROI's EL
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) [ r—

e [1343 | [a0sen [037e175 | (016174082 Relenaz6
Thaas 63 [sa11ss | [o1333s [a15  |[oraosos)  [33 | [s40 | | [m [ [ Hol ol
Thoas (825
Fazus ooy (o152 | [oooras [oossars | [oosssesze]  [ssos|[s;iz_| [s;13|[ssm | [s3|[sa
Tham (67 [s2ams | [noz7sa [a022s | [oommossos]  [e [z | (36 (a4 | [a5s|[3m0
Re225 [125) [43825 | [o.107a5 [oweizs | jpimaoz  (se0 | [sem s |[woi | [100s | [1o12
Fo2ta (242) [Samses |[o7s® | [0 (o |  Remmazzs
Po2ne 295 [frev7s | [z 55z | [ararem s [
Pozis @2l [rossize |[omsos | [1aizs | [orrssae e
6214 (e09) [ [0 1513278
N [sawm (o195 o 1epzee .
o228 (338) (0702 | [025383 [03%2s | [orarorsz]  [17m0 | [1795
hozz (811) e | [ozaz [navizs | [ooszooo;n,  [a7es | [s800
o220 (569 [30%67 | 012815 [m1g7s | jporomssz  (me | [s11s
et [25282 | 013254 [o [o [1288 | [125%
i [arims | [1asses | (206w | joamemimz  [1zas |1
e e [zz8% | [oosar7 [ooieiz | jposesizz (2| [zi6 ] ]
P (75015 | (o320 [oesesrs | [oizsoszz  [sovd | |30es ] ]
a0 (s (130204 | (oo [o4zs625 | [oosrzeres] 77w | [zraz | [z7a3_|[r7er | |[7ee |[77e |
e (1700252 | [0:20095 [o00125 | [oo77aroes|  [aam0 | (309 | (30 || | [ass0 | (360
87 el [a3708 | [0zt [o070125 | [oo720mo67| 516 | [2529 | [25%0 |[z5%6 | [a616 | (2620
Mt (35) o o [oosters | [ooseororzl [ o | C—— 0
oz ) |00s77 | [000%s | mazze (eeszss) (03072 | [nonizz et
hoz2s ga/aim (00672 | (000456 | maos (ieesy (023943 | [nooiss B e
o2 s [0visse | [0007en | mazs (/03 [03ws | [ooomt
Commarts

190 103 - Updsied 13 Aprl 2010 WP Call - OK.
3553 101 | Conocied 242K6\ and updated 40K, and 137Cs salrations P, 21 Apr 2010, Call - OK.
€353 105 Updsted ros ftom 0 Goamaty, June 2010 UF). Caib OF





Detector G, Geometry R
30
[image: image20.jpg]B EfficiencyBackGround

Detector iG ~| Geometry R ~  Current Setting 103
Efficiency Background ROI's EL
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) [ r—

e [135512 | [007346 [037e175 | (016174082 Relenaz6

Thaas 63 [1s7e1 | [o0usss [a15  |[oraosos)  [33 | [s40 | | [m [ [ Hol ol
Thoas (825

Fazss oo [00med | [000aTs [oossars | [oosssesze]  [ssos|[s;iz_| [s;13|[ssm | [s3|[sa

Tham (67 [oz2s1 | [ooiass [a022s | [oommossos]  [e [z | (36 (a4 | [a5s|[3m0

Re225 [125) [tes72s | [o0a0n7 [oweizs | jpimaoz  (se0 | [sem s |[woi | [100s | [1o12

Po2ts (202) e | [0 o | rewmez

Po2ne 295 e | T s [

Poz1s (52) [oziess | [1a12s | [ox77aaze  [es7

6214 (e09) [0 1513278

N [Gomer (575500 .

o228 (338) (2770 | [o1118 [03%2s | [orarorsz]  [17m0 | [1795

hozz (811) (250652 | (010415 [navizs | [ooszooo;n,  [a7es | [s800

et (145854 | (005068 [m1g7s | jporomssz  (me | [s11s

et [13201 | [oosizs [o [o [1288 | [125%

i [avesi7 | jomeass | (206w | joamamimz  1zas |1

e e [ozs06s | [oonasz [ooieiz | jposesizz (2| [zi6 ] ]

P [7arzi | [a1soen [oesesrs | [oizsoszz  [sovd | |30es ] ]

ko (e [onme | [ooms [o4zs625 | [oosrzeres] 77w | [zraz | [z7a3_|[r7er | |[7ee |[77e |

e (21323 | [03305 [o00125 | [oo77aroes|  [aam0 | (309 | (30 || | [ass0 | (360

87 el (22| [ozan2 [o070125 | [oo720mo67| 516 | [2529 | [25%0 |[z5%6 | [a616 | (2620

Mt (35) o o [oosters | [ooseororzl [ o | C—— 0

ac2z6 e [003a4 | (00053 | Re2z6 (186/295 (03045 | [000267 et
hozzs Ga/am (004015 | (000572 | maos (iesssy (01881 | [noots, B e
o2 s [oveszz |[noowe | msze (w03 [0zrvz | [ooozse
Commerts

Updated by AE to mclude Fa atis. JP 1o update laer.
350102 Updsied 15 Api UF) o O
€353 103 UpGated ros fom G aecmaty, Juns 2010 UP). Callb OK.





Detector G, Geometry V
31
[image: image21.jpg]B EfficiencyBackGround

Detector iG ~| Geometry |\y ~  Current Setting 1
Etficiency Background ROI's o Tl
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)

e [a7os3e | [a10385 [037e175 | (016174082 Relenaz6

Thaas 63 [7o7ese | 018077 [a15  |[oraosos)  [33 | [s40 | | [m [ [ Hol ol
Thoas (825

Fazus ooy [pasiz | [oonios [oossars | [oosssesze]  [ssos|[s;iz_| [s;13|[ssm | [s3|[sa

Tham (67 [asies1 | [o0se7s [a022s | [oommossos]  [e [z | (36 (a4 | [a5s|[3m0

Re225 [125) [1sezes | [orz2nt [oweizs | jpimaoz  (se0 | [sem s |[woi | [100s | [1o12

Fo2ta (242) [z [otssss | [0 1o |  Rewns2s

Po2ne 295 [ievaez | [a5000 55z | [ararem s [

Poz1s (52) mows | [oszr | (11125 | [oir7sane  [es7

6214 (e09) [0 1513278

N 575505 (575500 .

Ac228 (3381 0201 (03925 | [orermiszz 170 | (1795

hozz (811) 020578 [navizs | [ooszooo;n,  [a7es | [s800

et 013681 [m1g7s | jporomssz  (me | [s11s

e (riraz | [o1eem o o (245 ] 175

i [rizoosz | [1esose | (2067 | joamamimz  [1zas | [12

e e [o5ses | [o0zres [ooieiz | jposesizz (2| [zi6 ] ]

P [5527 | [oameaz [oesesrs | [oizsoszz  [sovd | |30es ] ]

kio (e (L7106 | [00eoes [o4zs625 | [oosrzeres] 77w | [zraz | [z7a3_|[r7er | |[7ee |[77e |

e Zssema | (421402 [o00125 | [oo77aroes|  [aam0 | (309 | (30 || | [ass0 | (360

87 el [a57m | [07ss8 [o070125 | [oo720mo67| 516 | [2529 | [25%0 |[z5%6 | [a616 | (2620

Mot (@39) [zzaz_|[s73 [oosiers |[ooseororzl [0 o | [0 _Jlo_| [0 |0

oz ) |00703 | [0m2ee | mazze (eezss (030313 | [oo0ss et
hozzs (sa/e1n) (008438 | [0015% | maos qiesssy (01301 | [Donzst ==
o2 s (01 | [oozs | mazs (s [03zeis | [ooosit
Commerts

Updated by AE to mclud Fa atis. JF 1o updato laer
352100 Updsiod calbrations for 3 . Swoept 750 and SAMn 21 May 2010 WP Cal - OK.
353101 Updatod i with goom 3 11 Juns 2010 0P} Call- G





Detector N, Geometry A
32
[image: image22.jpg]B EfficiencyBackGround

Detector }N > Geometry A ~ Current Setting 11n.105
ate [10/08/2010 23253 P _|
Efficiency Background ROI's P ezmozmid f
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backarourc)
e (205755 | [o0mez [oBsrasioe] [0aziaese]  Reieinazze
Tho 63 [s0716s | [0.10056 [as7ozs2se] [oseooere] 337 | [343 S [ | [mr ][ 3 Horol
Thoas (825

Pazis ooy (013 [[oonas | [oosesoses|[oomesose]  [seso |[sao0 | (s |[s0 |

Thawm (87) [oas21z | o0zt [oomuzsse [ooessezz: (32 (s | (w8 |l |

Raz2s (156) [25229 | [ocmsez | [ossvasise [oazi7asse (s |[sss | (s |[10o7 |

Po2e 2621 [ameis Jloroe | o o | ;eweess

Fe21e om) [seni7r | jozerr | |oessa osevedss (1w | [15e8

Fo214 (3521 [Toos1es | [oamer | |17z [osesmost (1055 | [1om5

T s —
e [eaosr [11aae511¢ 030453031

rozze 81 (57300 | [0zoes | [ozis0eone [ooeeeze 1786

s [asuss | 010450 [oresosore [0szzses  [s777 (061

pocs (s [2eews |[owssr | [ossassue [orresme (oo [s1es

Paca ay [24sns | [ovoer o o [1218 (1312

ooz 3y |zmroes | (vioser | [vosiesred] s [12e || e iz

gt g [oomer | [oomoe | [oossoas [orrasere

oo ey [roeras |[oseos | [odoassars| ossrizees]  [soes

w0 (e [ [cos [o0sstesc (01320351 [r70m

e 2177166 | (067201 [008053022| [0 07581657 [es3

b7 w3198 | oows 006790657 [0ossonss| (2528

oSt @ (1708 | oosis 07608611 [012msser  [s3e0

o7z (ewaw) (0045 | (00008 | Razas (eess)  [0zo611s | (0004306 Pasemord
ho228 (ga/an) (005256 | (000443 | Rapos presseszy (0163995 | [0002591 [ervadorrad
o2zn s [oossre | (000757 | Razes (eson)  [ozrasrs | [somets
Comments

Calln 103 - Coneced 185keY calbestions 15 Api 2010 UF Orginal 137G & 40K. Calb - PR
ESNIn 101 | updaied 1372, 25 Hay 2010 UP) Calb- O
a1 105 Updated ro, Juno 2010 UF). Cal - OK.





 Detector N, Geometry B 
33
[image: image23.jpg]B EfficiencyBackGround

Detector }N & Geometry [ o Current Setting I 103
. ate [21/05720103 1524 PV |
Efficiency Background ROI's fer PUTPRIETEETIIN
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)
e [1.5755 | [a0ss0s [oBsrasioe] [0aziaese]  Reieinazze
oo 631 (382875 | [003672 [os70zszse] [ozseosere] [z | 340 e | w2 [ HolRol
Thoas (825
Fazus ooy (013 | [oooere [oosseosee| [oostesose]  [ses0 | [saon | [saon |[s=0 |

Thawm (87) [oas257 | 00208 [oomuzsse [ooessezz: (32 (s | (w8 |l |

Raz2s (156) [oosote | (010503 | [ossrastse [pazi7asse  (sso |[sss | s |[10o7 |

Po2e 2621 [ooes Jfovos | o o | ;ewmezs

Po2ne 295 [sss | jozst el = =

Po2is (352 70e7 17z [oooomonc]  [1ess | [1o05

i 3 vaassire| [0 masann

rozze 81 [ oanzs | [ozsoesre| [oaeazel  [17es

s (6213 | 01680 [oresosore [0szzses  [s777 (061
pocs (@ [27s |[ovoos | [oosassue [orresme (oo [s1es
Pace ey [zsost|[oosere | o o [1218 (1312
B I T2 O Sz X7 R I e iz
gt g [osse7s | [oomrs | [ooessoas [orraseres

oo sy [Pire |[osesss | [odvessars| oswrizees]  [soes

w0 (e [t |[oowss | [ocssveed| [ovaeasm|  [rroe

e 210024 | 065373 [008053022| [0 07581657 [es3

b7 w3108 | oows 006790657 [0ossonss| (2528

oSt @ (1708 | oosis 07608611 [012msser  [s3e0

o7z (Gvaw) (004 | (00040 | Razze (oss)  [0a02ss | [0mast Pasemord
Ao22s (ga/a1n) (00508 (000486 | mapos piessssy (01636 | [000859 [ervadorrad
o2zn s [0osetr | (000 | Razes (eson  [ozress | [soreis
Commerts

[Dodated b AE to nckids Haratos. JP updatod 511 & Th serios ratos and affiioncies, 27/Aua 2008 Version 1015 coectod erto in enty of 40KeV ifciency.
C2n 103" Uipdsiod all bu 10%. 14 Apri 2010 UF) 3t OK.
€2 103  Usasied 0K, 21 M3y 2010 UF). Calb- OF





Detector N, Geometry Q

34
[image: image24.jpg]B EfficiencyBackGround

Detector }N ~| Geometry [ ~  Current Setting n.104
Etficiency Background ROI's o PITZIEETIE
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)
e [i6s16 | [oomi7s [oBsrasioe] [0aziaese]  Reieinazze
Thaas 63 77652 | 012208 [as7ozs2se] [oseooere] 337 | [343 S |[me | [ [ [ Hol ol
Thoas (825

Pazis ooy (01we |[ooosr | [oosesoses|[oomesose]  [seso |[sa0 | (s |[s0 |

Thawm 67) [o5e7a5 | [o.03391 [omizzosse [oveseszsz [z [ | s | @7 |

Ra225 [155) (15504 [o:10883 [ssrastee] [oazivasss] [0 | [ses i

PB-214 (242) [524283 |[o130e3 | [0 |0 | Refer Ra224 -

Pozis %) [treisse |[ozrie |  [osezee omeowesy]  [iess | [1oes

rozie g [teze |[osams | [1reseese| oseseoec]  [tess | [1aes

ro10 19 [orzms | [rastesrae] [o3mdssos

. oz | oz oswer| 17

s sas7 | 023 [o1ss0697e| [ou3zase|  [a777 (51
o 3250z | 013214 [oosaesnie| [o1772801]  [soes (516
ozt f241) [seuss | [o1s3s o o [1245 52
i [szresss | 1miza | [iosier pasois  12s EREEN
e e [ | [oosaia [comsea52 [a2037ee

T (g (ose | caower s e

Yo e e oo | osses orsee e

e 210054 | (064734 [0 06033027 [0 07581657 =

Hlhe [e3138 | [o116e [o06790557| [o076sm0es] 2502

Masa (838) o o [o160sss11| (0128085 4380

hozza (evaw) (0059 | (000K | Razas (esss)  [no67ad | (00035 rasomord
Ao22s (ga/an) (005028 | (000605 | mapos (iessssy (02268 | [000228 [ervaderrad
fno22s (sasss) [o1022 | 00102 | Re2s issvevs)  [ossors | [o00sss
Comers

Usdated AFE and TF with eficencies provided by JF (U and Th standards] 16/11/2008. How version s Cal9 102
Ca5n 103  Updsiod all but 01 nods hor chocking. 14 Aprl 2010 0F] Call - OF.
35101 - Updsted 108, 21 May 2010 (P} Calb - OFf





Detector N, Geometry R
35
[image: image25.jpg]B EfficiencyBackGround

Detector }N 3 Geometry R o Current Setting n.104
Etficiency Background ROI's Dee [1orme/zcto 2z M |
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backarourc)

uzs s [135652 | [o0mzs [oBsrasioe] [0aziaese]  Reieinazze

Thow (63 [216271 | [oosaa [as7ozs2se] [oseooere] 337 | [343 S |[me | [ [ O3 Howriol
Thoas (825

Fazus ooy [oowes | [o0ses [oosseosee| [oostesose]  [ses0 | [saon | [saon |[s=0 |

Tham (67 27| ooz [norzszose [oossesess (s || | |me || |

Re225 [125) [ts1s%2 | 004635 [ass7atee| [oazi7asse] [0 | [ses s | 007 |

Fo2ta (242) [zs0ses |[oosese | [0 |[o |  Pewmazzs

Fo2ts 2 (06321 | 014635 [oezzmzone josenadeer,  [1%e | 1568

Po2is (352 s |[oaer | [1arzseese| [osmmone] [1ess | [1oes

N o |[ooeoz | [ 1asssrie (0w

Ac228 (3381 (33225 | 013308 [ozis0657¢| [03s6an70e,  [1786

ho228 @11) (27383 | [o1m0 [oros0697€| [03azeasee|  [s727 [aes1

Ao (563 190836 | (007643 [nosagsnie (0177208 [soes (5166

Raz24 (2a1) [0 | (00032 o o [12a8 [1312

har ) [ieo0zes | [oeanes | [10siez zadsmisin 1285 | « | i3z

Foziz 40 (034028 | [oonaan [o04s3022) 011202758

e (a3 [a57602 | [a15328 [odvéeearz) [pesvizeez  [068

ka0 aen (087 | [oozrs [od0saneec| (01326361 [770m

e (1323847 | 039042 [ooeosonzz| [oo7smmes:) [

Bo7 (8 [25385__| 019907 005750697 [007ses0a]  [2524

Mot (@39) [o o [o160s6511) (012626597 [a380

oz aae)  |0bans | [o0meT | mazze (eess)  [0aoss | [ooos et
Ac22s ey (00002 [[00018 | Razes sssssz) (01917 | [oom7z [ervadorrad
oz s [0070n | [ooome | maze (g3  [02nss | [omozs
Commarts

Fio Cals 102 Undated 15 Mar 2010 by JP verfied by AFB.Col - OK.
351103  Updsied uncert on ratos 15 Apii 2010 (P), Calb - OF:
35101 Updated ro flom G asoréary, Jina 2010 P, Caib - O





Detector N, Geometry V
36
[image: image26.jpg]B EfficiencyBackGround

B 02513 AM |

Detector |y ~| Geometry [y ~| cCurrent Setting 102
Efficiency Background ROI's g
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)

uzs s (72106 | [ata86n [oBsrasioe] [0aziaese]  Reieinazze
Tho 63 [ 76824 | 010455 [as7ozs2se] [oseooere] 337 | [343 S |[me | [ [
Thoas (825
Fazss ooy [oi7esz | [oooez [oosseosee| [oostesose]  [ses0 | [saon | [saon |[s=0 |
Tham (67 [sesss | (003283 [norzszose [oossesess (s || | |me || |
Re225 [125) [3s90a2 | [a10000 [ass7atee| [oazi7asse] [0 | [ses s | 007 |
Fo2ta (242) [Goss |[otsoe | [0 (o |  Remmazzs
Po2ne 295 [Ta7065 | (35522 el = =
Pozis @2 [zievses |[osts | [1ivzeese) [oemoer]  [1ess | [1ess
6214 (e09) [ == T
Po210 46 [5rs|  [trasssrie][paaie
Ac228 (3381 (500356 | (024537 [ozis0657¢| [03s6an70e,  [1786
ho228 @11) a7 | (0222 [oros0697€| [03azeasee|  [s727 [aes1
Ao (563 [amzrs | 013028 [nosagsnie (0177208 [soes (5166
Raz24 (2a1) [sas74 | 014258 o o [12a8 [1312
har ) [pez2s | [1amaos | [1osiezt [zedsoisn 1285 | « | i3z
Foziz 40 [aazr7a | [o02er7 [o04s3022) 011202758
e (a3 [sa4z43 | o390 [odvéeearz) [pesvizeez  [068
ka0 aen (S| (o207 [od0saneec| (01326361 [770m
e (31267 | [pz17z8 [ooeosonzz| [oo7smmes:) [
Bo7 (8 [T [1 005750697 [007ses0a]  [2524
Mnss (835) [ [T [o160s6511) (0126268597 [a380
oz ae)  [006zee | (00079 | mazze eeszss) (0301 | [oouser
oz e (007027 | 00085 | mazss qesssy (019106 | |D0S
oz s [01isw |[omiszs | maze (es0s) [0z | [oooses

Comments

a0 100 - Updsied FOIs (om0} JF 15 Apr 2010 NOT CALIBRATED
350101 - Updstod all Ines salbratons, stcept 75 and 53, 21 May 2010 P, Caib - OK
5102 caected S38he Ins caib, 25 May 2010 (). Calls OK.

1 HoFol

Password
[ervadervact





Detector O, Geometry A
37
[image: image27.jpg]EfficiencyBackGround

1 HoFol

Password
[orvadersad

Detector io ~| Geometry A ~  Current Setting 1

Efficiency Background ROI's e ove oo azizan |

(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)
e (221356 | [a07a6n [oa15z507] [021411752]  Reieema2z6
Thaas 63 [45403 | [o1178 [oavissone] [or7mazre (s | [349 ®0 |3 | [@ [
Thas @25
Fazus ooy (o | [ooniez [oueis] [0osrsans  [san | [saes | [sam | (s | [swz |0
Thza0 67 [os0r2z | o0t [ozsoeores [pzvazsisn] (s || | (w |[s | e [
Raz2s (156) [176s | [oorees [oaiszeore [pavarizez  (ses | [ssm o |07 | [ | (10
Po214 (242) [a7eet1 J[oizisa | [0 |0 | ReferRa22e -
Foz1s (239) [10z120: | 025322 [odsamiere] [oieaosmsd  [1ss | [1578 | [187a | [1ss7 | [16m |[1638
Poos 52 = [aerzsomes| o 1stocs oo e e s e
6214 (e09) [1052345 | (025785 E e
T [ses0rs | 012582 [0 | [z |z | |za_|[zs
o228 (338) [s09420 | [0z0229 072620527 [0796963e2) 1807 | (1005 | (1605 | [16z6 | |[t039 | [16a7
hozz (811) (w3077 | (0178 [o71zo076¢| [po77obazs|  [am0a | [sms | [seem | [asz3 | [abva | lsns
et 262668 | 010583 [oosnisx poseaies  ma | (i3 | s | me0 | me | s
et [zass | o108 o o =) [1316 | [1323
i [zaz7ss | [1i7azs | [oeeaorew jnzmesrrie  [1ms | [12e8 [16 | [1323
e e [T-0295 | [o0waaz [oosisz0] [oizeevz, [z | [z D
P [ses215 | [nzrem [oziasiere [oriazsooc [0z | |08z INEEmIET
ko (e [7s | (oo [oasazeozs [ooreazan  |7rs | [77es | | [rrer | [y
e (21276 | [osor7 [o0sesonse (oosaoroes  [asz | 352z | (365 | (376
87 el (37055 | (00345 [o0ss6se1) [ori07i00e, (2537|2598 | (w10 (230
Mot (@39) [17006 | [ooms [oo7osese) (00sezzrar  [aa0s | [saza | |sazs | |asr2 | |asr2 | |sss2
oz ) (005197 | (00097 | mazze (eeszss)  [0asee | [ousy
pozze ooty [00553% | (000388 | mazes risersz) (019087 | (000214
o2 s (0100005 | [oneses | mazze (16s/E03)  [0300s6 | 00034

Comments

Updated 12 Ao 2010 JP. valdaled AF8 - Coll - OK.
310101 - Updsiod atio unoaie 15 Api 2010 WP Call - OK.
Calla 102 Updsied o uoing geom O 10 June 2016 (UF). Caib - OK





Detector O, Geometry B
38
[image: image28.jpg]EfficiencyBackGround

Detector io ~| Geometry g ~  Current Setting 120.105
Efficiency Background ROI's e sz austen
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)

e (211757 | [aorsie [oa15z507] [021411752]  Reieema2z6

Thaas 63 [r0s1 | [01021 [oavissone] [or7mazre (s | [349 ®0 |3 | [@ [ 3 Horol
Thas @25

Pazss ooy (o1 | [poore [oueis] [0osrsans  [san | [saes | [sam | (s | [swz |0

Thza0 67 [oase72 | [o02ees [ozsoeores [pzvazsisn] (s || | (w |[s | e [

Raz2s (156) [2562% | 008438 [oaiszeore [pavarizez  (ses | [ssm o |07 | [ | (10

o214 221 (w52 (o127 | [0 |5 |  Rewns2s -

Foz1s (239) [srioze | jozmn 04638461z 016903554 s | 178 | [ |[r | [1e: | [1e:

Poos s e o | omze 01e e ] [ o] e [

P [ [ro ] ] B ] [

o228 (338) [a5705 | [020247 072620527 [0796963e2) 1807 | (1005 | (1605 | [16z6 | |[t039 | [16a7

hozz (811) [37372 | 017798 [o71zo076¢| [po77obazs|  [am0a | [sms | [seem | [asz3 | [abva | lsns

et [25676¢ | (01058 [oosnisx poseaies  ma | (i3 | s | me0 | me | s

et [25320z | 010877 o o =) [1316 | [1323

i [ozer | [11sess | [oeeaorex nzmesrrie  |1ms | [12e8 [16 | [1323

e e [oeeare | [oowaes [oosisz0] [oizeevz, [z | [z D

P [seeezi | [ozrees [oziasiere [oriazsooc [0z | |08z INEEmIET

ko (e (100 | [ooes [oasazeozs [ooreazan  |7rs | [77es | | [rrer | [y

e EETTmi [o0sesonse (oosaoroes  [asz | 352z | (365 | (376

bo7  wE o [o [o0ss6se1) [ori07i00e, (2537|2598 | (w10 (230

Mot (@39) o o [oo7osese) (00sezzrar  [aa0s | [saza | |sazs | |asr2 | |asr2 | |sss2

oz ) |00z | (0002 | mazze (eeszss)  [0aem | [oomaz et
pozes (savern)  [005463 | (000809 | maces (imses (07672 | (000606 e
Aoz s (oo |[0mi3 | maze (g3  [0zws | [oorior
Commarts

Updated 137C 12 Apr 2010 P valdatod AFB. Call - OK.
320104 " Updsiod 15 Apr 2010 () Call - 0K
C322105  Updstod 40K lne, 15 A0 2010 UF). Caib - DK





Detector O, Geometry Q
39
[image: image29.jpg]B EfficiencyBackGround

Detector io 3 Geometry & o Current Setting 104
Etficiency Background ROI's e oo s=ooru
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backaroure)

uzs s [1e75es | [o0sems [oa15z507] [021411752]  Reieema2z6

Tho 63 [ss1662 | 013705 [asissese] [o17era2re]  [344 | [34s E e N | E [ HolRol
Thoas (825

Fazss ooy [oissz | [oooves [ooseise| [oosoreaa) s |[sazs | [s3 |[ssa7 | ([s3vz_|[saen

Tham (67 [ns0z1e | 002935 [ozsoeores [pzvazsisn] (s || | (w |[s | e [

Re225 [125) [as130s | 011002 [aatszsore| [ozrarizes]  [s89 | [see o |07 | [ | (10

o214 221 [so572¢ | (01362 | [oesa0rew: [0zoeerrie  ReeRazze

Foz1s (239) [Tre7e | [ozsem | [ndeameerc (oteswmsd  [1ses | 1578 | [1s7s | [1ss7 | [ |[1eas |

rozis @2 [rosses |[oseerz | [srzeome] [ise | [1s75 | [rors 1o | [1see ] [1s1z |

Po210 46 [ovse [resmme]

Ac228 (3381 026143 072626527 [0736963e2, 1807 | 1008

ho228 @11) a3 | [0z2678 [or1zo076e| [poz7baze)  [ana | [smns (005

Ao (563 [s2s82 | 013325 [nosnisx poseaies  ma | [si 5153

Raz24 (2a1) [a7e7s | [p1saes o o [1ss | 1268 [1323

har ) [sedmzss | [1assm | (oeeaorex: namesrrie  |1ms |18 [

Foziz 40 [3022 | [ooo77e [oosisz0] [oizeevz, [z | [z

e (a3 [51azs | 03020 [oziasiere [oriazsooc [0z | |08z

ka0 aen  [15ar_| [o0sote [oasazeozs [ooreazan  |7rs | [77es | 777 | [77a7

e Zapsen | (076276 [o0essse| [ooseo7oer| (12| (3522 3565 | (3578

Bo7 (8 (722300620 [o0ss6se1) [ori07i00e, (2537|2598 (2610 | (2630

Mot (@39) [o o [n07osssen) (00sezzrar  [as0s | [sazs | [a572 | 3502

oz ) |0ve | (00046, | mazze (eezss  [0aee2 | 000wt et
ez e (0062 | (00059 | mazse qesssy 02373 | Doz ==
Aoz s (010769 | (000923 | mazze (19603 (033 | 00034
Commarts

Fio CAL80 102 25 & o cvpeimented it including backorounds for 241 and 242 peaks. WP] Cal- 0K
750,103 Undaied 15 Apil 2010 1) Chock 40K Caib OF.
€350 101 - Updatod 40K and 137Cs caitiaions. Caib - OF





Detector O, Geometry R
40
[image: image30.jpg]EfficiencyBackGround

Detector io ~| Geometry R ~  Current Setting 160.1
Efficiency Background ROI's e ove a0 2zama P
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backarounc)

e [1z24758 | [a03477 [oa15z507] [021411752]  Reieema2z6

Thaas 63 [233m15 | 003 [oavissone] [or7mazre (s | [349 ®0 |3 | [@ [ 3 Horol
Thas @25

Fazus ooy [oosoas | 000405 [oueis] [0osrsans  [san | [saes | [sam | (s | [swz |0

Thza0 67 [ozstee | [o0r5as [ozsoeores [pzvazsisn] (s || | (w |[s | e [

Raz2s (156) [19575 | [oosar [oaiszeore [pavarizez  (ses | [ssm o |07 | [ | (10

o214 221 [zs0s [oomios | [0 (o |  FewnAszs -

Foz1s (239) (22165 | (015038 [odsamiere] [oieaosmsd  [1ss | [1578 | [187a | [1ss7 | [16m |[1638

Fo214 (352) [1027652 | (024765 | [oezaoves| [0 151608 % |[1o78_| [1e7s_| 1w | [1%0z | 1912 _

T [0 | [z |z | |za_|[zs

o228 (338) (330455 | 019385 072620527 [0796963e2) 1807 | (1005 | (1605 | [16z6 | |[t039 | [16a7

hozz (811) (19297 | 012582 [o71zo076¢| [po77obazs|  [am0a | [sms | [seem | [asz3 | [abva | lsns

et [Lonzsz | [ooras [oosnisx poseaies  ma | (i3 | s | me0 | me | s

et [16ss1e | 006623 o o =) [1316 | [1323

i [i6seme | joessze | [oeeaorex nzmesrrie  |1zs | 128 [16 | [1323

e e [o3843 | [oorses [oostsza07 [pizemvzw [z D

P [d.aam5 | [o17008 [ozzsiere [oriazsmor,  [s0e2 INEEmIET

ko (e |08z | [oomas [oasaeez [ooraazan|  [77s | [rrer | [y

e [1z6mz | [oa0n1z [o0sesonse) [ooeoorees  [3s1z | (365 | (376

87 el [543 | [oz09m [o0sss6sen) (oriomone (2537 | (w10 (230

Mot (@39) o o [o07osese) (00sezzrar  [aa0s | [saza | |sazs | |asr2 | |asr2 | |sss2

oz Gwae)  [00aa7e | (0002w | mazze eerzss) (03009 | o7 et
pozes ooy [004763 | (000262 | maces qimssee  [01967 | o008 e
o2 s [oveziz | [oooase | mazs (19603  [0zee25 | [ooozer
Commarts

Flo CALGD 105 Updated K40 and C=137 uncert Cal- 0K
350 107 - Updstod 15 Api 2010 (1P] Calb 0K
€352 103 Up3ated ADIs iom geom 0. June 2010 WP Cal- DK





Detector O, Geometry V
41
[image: image31.jpg]B EfficiencyBackGround

Detector io 3 Geometry [ - Current Setting 1
Etficiency Background ROI's b P
(kiee/ea) (s erorzs)  (etkses/oa) (b sner2s)  (Lstackoun)  (Pask Courts) (Rl Backarourc)

e [as6s¢ | [012402 [oa15z507] [021411752]  Reieema2z6

Thaas 63 716886 | [o17827 [asissese] [o17era2re]  [344 | [34s ®0 |3 | [@ [ [ Hol ol
Thoas (825

Fazus ooy [oiszs | (oo [ooseise| [oosoreaa) s |[sazs | [s3 |[ssa7 | ([s3vz_|[saen

Tham (67 [a79574 | [noz0mn [ozsoeores [pzvazsisn] (s || | (w |[s | e [

Re225 [125) [aza05 | 012022 [aatszsore| [ozvarizes] (89 | [see o |07 | [ | (10

Fo2ta (242) [z1s3 foi7s | (o 1o |  Rewns2s

Fo2ts 2 [1532642 | [037028 [aaesesere| [oieoomss)  [1ses | [1578 | [1578 |[1s%7 | [1em | [1e38

Pozis @2 [zesoess |[osse | [srzeome] [ise | [1s75 | [rors 1o | [1see ] [1s1z |
e =
T [7res | o150 01532080

o228 (338) 555102 | (026524 072626527 [0736963e2, 1807 | 1008

hozz (811) 6257 | [0z2607 [on72s07ee| [oo77snaze| (a0 | [am1s [a0ms

o220 (569 [3298 | [o13209 [o0ersc| [oosesries] (21| 5133 5153

e [w0s7as | 01638 o o [12s | 1268 ED

i [eseiaz | [1oseer | [oeeaorex jnzsesrrie  [1ms | [128 [

e e (05205 [o0zser [oosisz0] [oizeevz, [z | [z

P [seari | [a3re [oziasiere [oriazsooc [0z | |08z

io (s [16emee | (00603 [oasazeozs [ooreazan  |7rs | [77es | 777 | [77a7

e Za0msn | [7.60963 [o0essse| [ooseo7oer| (12| (3522 3565 | (3578

Bo7 (8 0 1 [o09830461| (01107100 (2837 | 2548 (2610 | (2630

Mnss (835) [ [T [n07ossse) (00sezzrar  [aa0s | [sazs | [a572 | 3502

soz Gwme)  [00mss |[0m7oz | mazze (eeszss) (03 | [ooiose et
hoz2s (G/a1m) (0039 |[0019%5 | maos (iesss (019839 | [0oossr B e
oz Euses  [ovenss |[omawz | maze (s [osa | [oonz
Commarts

[ ot uee ot carated
C355.10 - Ericred ai saibrations. sxcspt 780 and S4kn, 24 May 2010 UP), Calb - OK.
C3552 101 Updated i fom 0 asematy. Juns 2010 UF) Calb - DX






Appendix 3  Plots of absolute efficiencies for Detectors G, N and O

Detector G, Geometry A

[image: image32.emf]Det - G, Geom - A

0

2

4

6

8

10

12

14

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector G, Geometry B

[image: image33.emf]Det - G, Geom - B

0

2

4

6

8

10

12

14

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector G, Geometry Q
[image: image34.emf]Det - G, Geom - Q

0

2

4

6

8

10

12

14

16

18

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector G, Geometry R

[image: image35.emf]Det - G, Geom - R

0

1

2

3

4

5

6

7

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector G, Geometry V

[image: image36.emf]Det - G, Geom - V

0

5

10

15

20

25

30

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector N, Geometry A

[image: image37.emf]Det N, Geom - A

0

2

4

6

8

10

12

14

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector N, Geometry B

[image: image38.emf]Det - N, Geom - B

0

2

4

6

8

10

12

14

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector N, Geometry Q

[image: image39.emf]Det - N, Geom - Q

0

2

4

6

8

10

12

14

16

18

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector N, Geometry R
[image: image40.emf]Det - N, Geom - R

0

1

2

3

4

5

6

7

8

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector N, Geometry V

[image: image41.emf]Det - N, Geom - V

0

2

4

6

8

10

12

14

16

18

20

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector O, Geometry A

[image: image42.emf]Det - O, Geom - A

0

2

4

6

8

10

12

14

16

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector O, Geometry B

[image: image43.emf]Det - O, Geom - B

0

2

4

6

8

10

12

14

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector O, Geometry Q

[image: image44.emf]Det - O, Geom - Q

0

2

4

6

8

10

12

14

16

18

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector O, Geometry R

[image: image45.emf]Det - O, Geom - R

0

1

2

3

4

5

6

7

8

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Detector O, Geometry V
[image: image46.emf]Det - O, Geom - V

0

5

10

15

20

25

0 200 400 600 800 1000 1200 1400 1600

Energy, keV

Absolute Efficiency, %


Appendix 4  Tables of absolute efficiency values for detectors G, N and O

Detector G, Geometry A, B & Q

50
	Det - G
	
	Geom - A
	Geom - A
	
	Geom - B
	Geom - B
	
	Geom - Q
	Geom - Q

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.36906
	0.00610
	
	0.35139
	0.00834
	
	0.33569
	0.00492

	Th-234
	63
	9.78298
	0.13031
	
	9.99460
	0.20214
	
	14.43063
	0.17782

	Pa-234
	1001
	1.37543
	0.05980
	
	1.32921
	0.03790
	
	1.64507
	0.04013

	Th-230
	68
	11.44211
	0.45664
	
	11.83883
	0.34333
	
	16.65582
	0.36725

	Ra-226
	186
	8.05946
	0.10086
	
	7.90786
	0.17021
	
	12.48457
	0.15113

	Pb-214
	242
	6.14021
	0.07842
	
	5.97987
	0.12032
	
	7.50783
	0.09175

	Pb-214
	295
	5.11412
	0.06293
	
	5.03912
	0.10102
	
	6.25460
	0.07560

	Pb-214
	352
	4.30437
	0.05243
	
	4.25408
	0.08520
	
	5.24754
	0.06324

	Bi-214
	609
	2.10991
	0.02591
	
	2.09796
	0.04205
	
	2.50263
	0.03024

	Pb-210
	46
	8.68988
	0.11422
	
	8.69755
	0.17563
	
	12.77264
	0.15684

	Ac-228
	338
	4.17079
	0.08384
	
	4.12513
	0.08337
	
	5.53247
	0.11133

	Ac-228
	911
	1.57248
	0.03160
	
	1.58180
	0.03195
	
	2.03117
	0.04087

	Ac-228
	969
	1.50913
	0.03044
	
	1.52365
	0.03099
	
	1.94761
	0.03937

	Ra-224
	241
	5.81515
	0.11787
	
	5.64790
	0.11625
	
	7.91946
	0.16085

	Pb-212
	239
	6.20771
	0.12425
	
	6.15323
	0.12325
	
	7.96183
	0.15939

	Bi-212
	40
	8.02577
	0.17339
	
	7.71466
	0.18172
	
	12.17189
	0.26322

	Tl-208
	583
	2.10298
	0.04220
	
	2.13512
	0.04299
	
	2.61973
	0.05262

	K-40
	1461
	
	
	
	1.06810
	0.01682
	
	1.30999
	0.01651

	Cs-137
	662
	2.58279
	0.04005
	
	2.22696
	0.03445
	
	2.10387
	0.01670


50
Detector G, Geometry R & V

51
	Det - G
	
	Geom - R
	Geom - R
	
	Geom - V
	Geom - V

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.20967
	0.00294
	
	0.66480
	0.00907

	Th-234
	63
	5.22030
	0.06661
	
	20.47718
	0.25169

	Pa-234
	1001
	0.87941
	0.02237
	
	2.20036
	0.05949

	Th-230
	68
	5.94932
	0.18047
	
	24.49370
	0.48972

	Ra-226
	186
	4.66184
	0.05649
	
	13.96028
	0.17175

	Pb-214
	242
	3.48583
	0.04292
	
	10.50576
	0.12840

	Pb-214
	295
	2.95124
	0.03575
	
	8.76953
	0.10602

	Pb-214
	352
	2.52253
	0.03043
	
	7.31712
	0.08820

	Bi-214
	609
	1.32613
	0.01604
	
	3.26626
	0.03950

	Pb-210
	46
	4.05344
	0.05132
	
	19.00732
	0.23294

	Ac-228
	338
	2.43536
	0.04908
	
	6.13764
	0.12325

	Ac-228
	911
	0.98738
	0.01988
	
	2.24064
	0.04500

	Ac-228
	969
	0.94336
	0.01908
	
	2.13547
	0.04302

	Ra-224
	241
	3.21384
	0.06585
	
	10.13159
	0.20430

	Pb-212
	239
	3.22664
	0.06461
	
	9.46113
	0.18933

	Bi-212
	40
	2.87008
	0.07388
	
	5.48017
	0.13693

	Tl-208
	583
	1.22860
	0.02469
	
	2.92875
	0.05874

	K-40
	1461
	0.67975
	0.01626
	
	1.62143
	0.02858

	Cs-137
	662
	1.31920
	0.01990
	
	3.00813
	0.24793


Detector N, Geometry A, B & Q

52
	Det - N
	
	Geom - A
	Geom - A
	
	Geom - B
	Geom - B
	
	Geom - Q
	Geom - Q

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.3575
	0.0073
	
	0.3400
	0.0083
	
	0.2915
	0.0063

	Th-234
	63
	10.8583
	0.1448
	
	10.2100
	0.1276
	
	12.7357
	0.1627

	Pa-234
	1001
	1.4469
	0.0618
	
	1.5052
	0.0365
	
	1.6456
	0.0521

	Th-230
	68
	11.7898
	0.4321
	
	11.8018
	0.2787
	
	15.1126
	0.4520

	Ra-226
	186
	8.2221
	0.1204
	
	8.1026
	0.1477
	
	11.6879
	0.1500

	Pb-214
	242
	6.2165
	0.0796
	
	6.0744
	0.0745
	
	7.2126
	0.0897

	Pb-214
	295
	5.3284
	0.0656
	
	5.1669
	0.0625
	
	6.0920
	0.0742

	Pb-214
	352
	4.5118
	0.0550
	
	4.4181
	0.0533
	
	5.1548
	0.0624

	Bi-214
	609
	2.3276
	0.0286
	
	2.2785
	0.0275
	
	2.6044
	0.0316

	Pb-210
	46
	10.3626
	0.1354
	
	9.0597
	0.1122
	
	11.5599
	0.1458

	Ac-228
	338
	4.3629
	0.0880
	
	4.3876
	0.0885
	
	5.3432
	0.1079

	Ac-228
	911
	1.7502
	0.0353
	
	1.7639
	0.0356
	
	2.1090
	0.0426

	Ac-228
	969
	1.6879
	0.0342
	
	1.7078
	0.0346
	
	2.0466
	0.0416

	Ra-224
	241
	6.0466
	0.1239
	
	5.6809
	0.1167
	
	8.8228
	0.1800

	Pb-212
	239
	6.3538
	0.1272
	
	6.3343
	0.1269
	
	7.5157
	0.1505

	Bi-212
	40
	9.2660
	0.2083
	
	8.2847
	0.1888
	
	11.3204
	0.2580

	Tl-208
	583
	2.3238
	0.0467
	
	2.3598
	0.0474
	
	2.7121
	0.0546

	K-40
	1461
	
	
	
	1.2382
	0.0192
	
	1.5427
	0.0357

	Cs-137
	662
	2.5614
	0.0395
	
	2.4718
	0.0385
	
	2.4712
	0.0381


Detector N, Geometry R & V

53
	Det - N
	
	Geom - R
	Geom - R
	
	Geom - V
	Geom - V

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.2026
	0.0028
	
	0.4774
	0.0095

	Th-234
	63
	5.8205
	0.0725
	
	15.3820
	0.1929

	Pa-234
	1001
	1.0523
	0.0208
	
	1.9048
	0.0583

	Th-230
	68
	6.5913
	0.1614
	
	17.3315
	0.4338

	Ra-226
	186
	5.3877
	0.0652
	
	11.2248
	0.1407

	Pb-214
	242
	3.8614
	0.0470
	
	8.3916
	0.1037

	Pb-214
	295
	3.2927
	0.0397
	
	7.1142
	0.0864

	Pb-214
	352
	2.8386
	0.0342
	
	6.0154
	0.0727

	Bi-214
	609
	1.5507
	0.0187
	
	2.9524
	0.0358

	Pb-210
	46
	4.6461
	0.0576
	
	14.7993
	0.1836

	Ac-228
	338
	2.9143
	0.0584
	
	5.3365
	0.1076

	Ac-228
	911
	1.2494
	0.0250
	
	2.1045
	0.0424

	Ac-228
	969
	1.2003
	0.0241
	
	2.0206
	0.0410

	Ra-224
	241
	3.6442
	0.0732
	
	8.4894
	0.1730

	Pb-212
	239
	3.6704
	0.0734
	
	7.8744
	0.1577

	Bi-212
	40
	3.3691
	0.0710
	
	4.2351
	0.1197

	Tl-208
	583
	1.5010
	0.0300
	
	2.7680
	0.0557

	K-40
	1461
	0.8259
	0.0130
	
	1.5145
	0.0572

	Cs-137
	662
	1.5570
	0.0234
	
	3.6620
	0.4834


Detector O, Geometry A, B & Q
54
	Det - O
	
	Geom - A
	Geom - A
	
	Geom - B
	Geom - B
	
	Geom - Q
	Geom - Q

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.38834
	0.00655
	
	0.37157
	0.00659
	
	0.32910
	0.00512

	Th-234
	63
	11.87757
	0.15706
	
	10.78802
	0.13681
	
	14.71099
	0.18274

	Pa-234
	1001
	1.56469
	0.06422
	
	1.45407
	0.04202
	
	1.67273
	0.04250

	Th-230
	68
	13.36576
	0.48257
	
	12.23260
	0.32910
	
	16.05772
	0.39140

	Ra-226
	186
	8.93643
	0.11185
	
	8.33350
	0.11868
	
	12.71301
	0.15474

	Pb-214
	242
	6.58518
	0.08389
	
	6.25801
	0.07751
	
	7.64617
	0.09372

	Pb-214
	295
	5.60044
	0.06876
	
	5.27330
	0.06400
	
	6.41787
	0.07766

	Pb-214
	352
	4.71732
	0.05738
	
	4.45894
	0.05386
	
	5.38912
	0.06499

	Bi-214
	609
	2.31336
	0.02835
	
	2.21073
	0.02680
	
	2.59621
	0.03140

	Pb-210
	46
	11.50083
	0.14795
	
	9.75901
	0.12213
	
	13.19971
	0.00413

	Ac-228
	338
	4.41603
	0.08872
	
	4.35995
	0.08880
	
	5.69479
	0.11463

	Ac-228
	911
	1.69113
	0.03397
	
	1.66936
	0.03397
	
	2.14976
	0.04327

	Ac-228
	969
	1.65326
	0.03331
	
	1.61487
	0.03327
	
	2.07190
	0.04189

	Ra-224
	241
	6.66358
	0.13482
	
	6.14666
	0.12958
	
	9.18799
	0.18630

	Pb-212
	239
	6.72880
	0.13466
	
	6.48784
	0.13011
	
	8.35880
	0.16734

	Bi-212
	40
	10.19262
	0.21502
	
	8.55209
	0.21763
	
	13.17269
	0.28598

	Tl-208
	583
	2.25974
	0.04533
	
	2.25712
	0.04568
	
	2.79084
	0.05607

	K-40
	1461
	
	
	
	1.03830
	0.00838
	
	1.45732
	0.02377

	Cs-137
	662
	2.60320
	0.04010
	
	2.51343
	0.03889
	
	2.92420
	0.04487


Detector O, Geometry R & V

55
	Det - O
	
	Geom - R
	Geom - R
	
	Geom - V
	Geom - V

	
	
	absolute
	Efficiency
	
	absolute
	Efficiency

	Isotope
	keV
	efficiency %
	uncert 1sd
	
	efficiency %
	uncert 1sd

	U-235
	186
	0.21894
	0.00304
	
	0.58543
	0.01088

	Th-234
	63
	6.23251
	0.07867
	
	19.11985
	0.23503

	Pa-234
	1001
	0.97483
	0.02181
	
	2.07615
	0.05417

	Th-230
	68
	7.06288
	0.20657
	
	21.21962
	0.40013

	Ra-226
	186
	5.61996
	0.06803
	
	13.58757
	0.16909

	Pb-214
	242
	3.99134
	0.04888
	
	9.85690
	0.12032

	Pb-214
	295
	3.37878
	0.04083
	
	8.32324
	0.10054

	Pb-214
	352
	2.88666
	0.03478
	
	6.96880
	0.08396

	Bi-214
	609
	1.51946
	0.01834
	
	3.18056
	0.03842

	Pb-210
	46
	4.94884
	0.06192
	
	18.30047
	0.22367

	Ac-228
	338
	2.92505
	0.05862
	
	5.74651
	0.11546

	Ac-228
	911
	1.19556
	0.02395
	
	2.14722
	0.04314

	Ac-228
	969
	1.14315
	0.02293
	
	2.07421
	0.04182

	Ra-224
	241
	4.11451
	0.08281
	
	9.84828
	0.19890

	Pb-212
	239
	3.74430
	0.07491
	
	8.90858
	0.17829

	Bi-212
	40
	3.60747
	0.07762
	
	5.15345
	0.12805

	Tl-208
	583
	1.45752
	0.02919
	
	2.83315
	0.05685

	K-40
	1461
	0.80760
	0.01487
	
	1.59780
	0.02862

	Cs-137
	662
	1.48603
	0.02242
	
	2.83317
	0.44763


Appendix 5  Listing of gamma spectroscopy standards used for calibration

An up-to-date listing of all gamma spectroscopy standards is maintained in the file ‘All-ERISS-standards.xls’ on Sharepoint directory :
SSDX > Equipment and Techniques > HPGe Gamma Spec > Calibrations > Standards Listing

J Pfitzner





October 2010





(Release status – unrestricted)





eriss HPGe detector calibration





internal report





576





Detector – prior to starting calibration, ensure the optimum performance and stability of the detector, its associated electronic components and background reduction housing.





Select the geometry(ies) to be calibrated.





Count the background of the matrix used in geometry(ies) for 3 days.





Count Uranium and Thorium series standards and 40K and 137Cs standards, including standards for ratios. Count time must be sufficient to ensure < 1% statistical uncertainty of isotope peak counts.





Analyse all standards using Visual Gamma software.





From Visual Gamma, export ‘P’ files, saving in directory named for that detector.





238U series





232Th series





40K and 137Cs





In each detector directory, open Excel workbooks in three categories for importing ‘P’ files & performing subsequent calibration calculations.





Calculate 226Ra ratios at 186keV.





Calculate 232Th ratios at 63keV.





Calculate calibration factors and uncertainties for each isotope line.





Subtract 226Ra contribution at 186kev & calculate 235U calibn factor & uncertainty.





Enter calibration factors and uncertainties in Visual Gamma section for that detector and geometry.





Validate data entered in Visual Gamma. Update IR pages and SSDX Sharepoint with new data.





Ensure all data from ‘P’ file is copied into the Excel spreadsheet and column are correctly aligned.





Name the spreadsheet with the count code for the standard or background counted, e.g. ZV003n.02.





Check that the Livetime from cell A7, is correctly calculated in cell F7, in units of kiloseconds. Cell E7 identifies these values as ‘Livetime, kilosecs’. Cells A7, E7 and F7 are highlighted in the same colour to clarify the link.





Delete the test source columns as they are not used during the calibration procedures.





Calculating Peak Area, counts per kilosecond:


Cell ranges E13-E31 and G13-G31 are divided by the Livetime in kiloseconds from cell F7 and placed into cell ranges I13-I31 and K13-K31, respectively. Cells I11 & I12 and K11 & K12 contain the headers for these columns - ‘counts/ksec’, ‘Peak’ and ‘1sd’, ‘uncert’, respectively.





This completes the initial spreadsheet operations and calculations common to all calibration and ratio calculations.





For the next steps in the calibration procedures, branch to the appropriate section :








Flow chart for 238U series and  235U calibration








Flow chart for 232Th series calibration








Flow chart for 40K and 137Cs calibrations





Flow chart for calculation of radium (226Ra) ratios 








Flow chart for calculation of radium (232Th) ratios 








Subtraction of Background, counts per kilosecond:


From the appropriate background for this standard, the background counts per kilosecond at each isotope line are subtracted from cells I13 – I31 and placed in cells J13 – J31. The associated backgrounds are calculated in cells L13 – L31, using formula ‘=SQRT((StdLx)2 +(BgdLx)2)’ .





238U series & 226Ra ratios





232Th series & ratios





40K & 137Cs calibrations





Live count time, ksec





Total Activity, Bq





Peak Area/ live count time





Peak Area less Background Cts/Ksec / total activity, Bq





Peak Area Cts/Ksec less Background Cts/Ksec





The three 226Ra ratios are of the counts/ksec, 


186keV:295keV, 186keV:352keV and 186keV:609keV.





Cell M18 is ratio of 186keV:295keV, formula is ‘= J17/J18’.


Cell N18 is uncertainty in 186keV:295keV ratio, 


formula is ‘=M18*SQRT((L17/J17)^2 + (L18/J18)^2)’.








Cell M19 is ratio of 186keV:352keV, formula is ‘= J17/J19’.


Cell N19 is uncertainty in 186keV:keV 295 ratio, 


formula is ‘=M19*SQRT((L17/J17)^2 + (L19/J19)^2)’.








Cell M20 is ratio of 186keV:609keV, formula is ‘= J17/J20’.


Cell N20 is uncertainty in 186keV:keV 609 ratio, 


formula is ‘=M20*SQRT((L17/J17)^2 + (L20/J20)^2)’.








The weighted average of several counts of the 186keV:295keV ratios are calculated in the ‘summary‘ worksheet in cell C6, using the following formula:


‘ = (W1M18/(W1N18)^2 + W2M18/(W2N18)^2)/


(1/(W1N18)^2 + 1/(W2N18)^2))’


where W1 and W2 are names of two worksheets of ratio counts and M18 and N18 are cells holding the 186keV:295keV ratios and uncertainties in those worksheets.





The 1 standard deviation (1 s.d.) uncertainty on the weighted average for 186keV:295keV in cell C8 uses the formula, ‘= 1/SQRT(1/(W1N18)^2 + 1/(W2N18)^2)’ in the same syntax as above.





In the ‘summary‘ worksheet, similar calculations are performed for 186keV:352keV at cell D7 and 186keV:609keV at cell E7 and their respective uncertainties in cells D8 and E8.





This completes the 226Ra ratio calculations.


These results can now be entered into the Visual Gamma calibration page.








Calculating Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond in cell ranges I13-I31 are divided by the Total Activity in Bq from cell F4 and placed into cell ranges M13-M31. 





Details of all eriss standards, including total activity, is maintained in the Excel spreadsheet ‘All eriss gamma standards.xls’, located at SSDX>Equipment and Techniques > HPGe Gamma Spec> Calibrations>Standards Listing.


The total activity of the 238U which is contained in the standard is entered into cells F4, in units of Bq. Cell E4 identifies these values as ‘Total activity 238U, Bq’. Cells E4 and F4 are highlighted in the same colour to clarify their relationship.


The 235U activity (natural activity ratio of 235U:238U of 0.046029 * 238U) entered into cell I4, with the text, ‘Total activity 235U, Bq’, in cell H4 identifying the isotope. Cells H4 and I4 are highlighted in the same colour to clarify their relationship.





Calculating Uncertainty on Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond uncertainties in cells K13-K31 are divided by the Total Activity in Bq from cell F4 and placed into cell range O13-O31.





Calculating the contribution of 226Ra to the 235U line at 186keV:


In the 238U standard worksheet, the counts per kilosecond from cells M18, M19 and M20 are multiplied by the ratio from the ‘summary’ sheet cells C6, C7 and C8, and the results are placed in cells Q18, Q19 and Q20, respectively. The uncertainties are placed in cells S18, S19 and S20.


The weighted average of these contributions, in counts per kilosecond, is calculated in cell Q22 and the associated uncertainty in cell S22.





The 238U series standards contains both 238U and 235U in their natural isotopic ratio. The only gamma emission line of 235U suitable for calibration is at 185.720keV. There is no gamma emission line of 238U which is suitable to use for calibration.


The 185.720keV line of 235U is overlaid with counts from the 186.211keV line of 226Ra, a decayed descendant of 238U. Isotope peaks which are in such close proximity are beyond the resolution capacity of HPGe detectors, so an additional manual procedure is required to separate the contribution of two isotopes at the 186keV region. 


Fig. 3 displays the section of an Excel spreadsheet containing the calculations for this procedure.








The weighted average of the following calibration factors are calculated in the ‘summary’ sheet:


U-235 @ 186keV in cell D15 from cells Q30, associated uncertainty in cell E15 from cells S30


Th-63 @ 63keV in cell D16 from cells M14, associated uncertainty in cell E16 from cells O14


Pa-234 @ 1001keV in cell D17 from cells M15, associated uncertainty in cell E17 from cells O15


Th-230 @ 68keV in cell D18 from cells M16, associated uncertainty in cell E18 from cells O16


Ra-226 @ 186keV in cell D19 from cells Q34, associated uncertainty in cell E19 from cells S34


Pb-214 @ 242keV in cell D20 from cells M18, associated uncertainty in cell E20 from cells O18


Pb-214 @ 295keV in cell D21 from cells M19, associated uncertainty in cell E21 from cells O19


Bi-214 @ 609keV in cell D22 from cells M20, associated uncertainty in cell E22 from cells O20


Pb-210 @ 46keV in cell D23 from cells M21, associated uncertainty in cell E23 from cells O21


See Fig. 4














Calculation of 226Ra calibration factor at 186keV:


In cell Q34, the calibration factor for 266Ra is calculated by dividing the counts per ksec from cell Q22 by the ‘Total activity 238U, Bq’ from cell F4. The associated uncertainty is calculated in cell S34 using the formula ‘= S22/F4’.





This completes the calculations for the 238U series calibration.


These results can now be entered into the Visual Gamma calibration page.





Calculating the calibration factor for 235U at 186keV:


In cell Q30, the calibration factor for 235U is calculated by dividing the counts per ksec from cell Q26 by the ‘Total activity 235U, Bq’ from cell H4. The associated uncertainty is calculated in cell S30 using the formula ‘= S26/H4’.





Subtracting the 226Ra contribution from the 186keV line:


In cell Q26, the weighted average of the 226Ra contribution in cell Q22 is subtracted from the combined 226Ra and 235U at 186keV in cell I17 to give the 235U at 186keV, all in units of counts per ksec. The associated uncertainties are calculated in cell S26, using the same units.





Details of all eriss standards, including total activity, is maintained in the Excel spreadsheet ‘All ERISS gamma standards.xls’, located at SSDX>Equipment and Techniques > HPGe Gamma Spec> Calibrations>Standards Listing.


The total activity of the 232Th which is contained in the standard is entered into cells F4, in units of Bq. Cell E4 identifies these values as ‘Total activity 232Th, Bq’. Cells E4 and F4 are highlighted in the same colour to clarify their relationship.





Calculating Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond in cell ranges J13 – J31 are divided by the Total Activity in Bq from cell F4 and placed into cell ranges M13 – M31.





Calculating Uncertainty on Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond uncertainties in cells L13 – L31 are divided by the Total Activity in Bq from cell F4 and placed into cell range O13–O31.








The weighted average of the following calibration factors are calculated in the ‘summary’ sheet:


228Ac @ 338keV in cell D7 from cells M22, associated uncertainty in cell E7 from cells O22


228Ac @ 911keV in cell D8 from cells M23, associated uncertainty in cell E8 from cells O23


228Ac @ 969keV in cell D9 from cells M24, associated uncertainty in cell E9 from cells O24


224Ra @ 241keV in cell D10 from cells M25, associated uncertainty in cell E10 from cells O25


212Pb @ 239keV in cell D11 from cells M26, associated uncertainty in cell E11 from cells O26


212Bi@ 40keV in cell D12 from cells M27, associated uncertainty in cell E12 from cells O27


208Tl @ 583keV in cell D13 from cells M28, associated uncertainty in cell E13 from cells O28


See Fig. 5





This completes the 232Th calibration factor calculations.


These results can now be entered into the Visual Gamma calibration page.





The three 232Th ratios are of the counts/ksec, 


63keV:338keV, 63keV:911keV and 63keV:969keV.





Cell R22 is ratio of 63keV:338keV, formula is ‘= J14/J22’.


Cell S22 is uncertainty in 63keV:338keV ratio, 


formula is ‘=R22*SQRT((L14/J14)^2 + (L22/J22)^2)’.








Cell R23 is ratio of 63keV:911keV, formula is ‘= J14/J23’.


Cell S23 is uncertainty in 63keV:911keV ratio, 


formula is ‘=R23*SQRT((L14/J14)^2 + (L23/J23)^2)’.








Cell R24 is ratio of 63keV:969keV, formula is ‘= J14/J24’.


Cell S24 is uncertainty in 63keV:969keV ratio, 


formula is ‘=R24*SQRT((L14/J14)^2 + (L24/J24)^2)’.








The weighted average of several counts of the 63keV:338keV ratios are calculated in the ‘summary‘ worksheet in cell C20, using the following formula:


‘ = (W1R22/(W1S22)^2 + W2R22/(W2S22)^2)/


(1/(W1S22)^2 + 1/(W2S22)^2))’


where W1 and W2 are names of two worksheets of ratio counts and R22 and S22 are cells holding the 63keV:338keV ratios and uncertainties in those worksheets.








The 1 standard deviation (1 s.d.) uncertainty on the weighted average for 63keV:338keV in cell E20 uses the formula, ‘= 1/SQRT(1/(W1S22)^2 + 1/(W2S22)^2)’ in the same syntax as above.








In the ‘summary‘ worksheet, similar calculations are performed for 63keV:911keV at cell C21 and 63keV:969keV at cell C22 and their respective uncertainties in cells E21 and E22.





This completes the 232Th ratio calculations.


These results can now be entered into the Visual Gamma calibration page.








Details of all eriss standards, including total activity, is maintained in the Excel spreadsheet ‘All ERISS gamma standards.xls’, located at SSDX>Equipment and Techniques > HPGe Gamma Spec> Calibrations>Standards Listing.


The total activity of the 40K or 137Cs which is contained in the standard is entered into cells F4, in units of Bq. Cell E4 identifies these values as ‘Total activity 40K (or 137Cs), Bq’. Cells E4 and F4 are highlighted in the same colour to clarify their relationship.





Calculating Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond  in cell ranges J29 - J30 are divided by the Total Activity in Bq from cell F4 and placed into cell ranges M29 - M30.





Calculating Uncertainty on Calibration Factor, counts per kilosecond per Bq:


Counts per kilosecond uncertainties in cells L29 - L30 are divided by the Total Activity in Bq from cell F4 and placed into cell range O29-O30.








The weighted average of the following calibration factors are calculated in the ‘summary’ sheet:


40K @ 1461keV in cell D6 from cells M29, associated uncertainty in cell E6 from cells O29


137Cs @ 661keV in cell D7 from cells M30, associated uncertainty in cell E7 from cells O30


See Fig. 6





This completes the 40K & 137Cs calibration factor calculations.


These results can now be entered into the Visual Gamma calibration page.





Start program Visual Gamma.





Click on the ‘View Detector Settings’ button.





Enter the password to allow data entry.








Select the correct detector and geometry from the drop-down menus.


Caution – be aware that movement of the mouse wheel may switch the display to another geometry without warning.








Enter the calibration factor and uncertainty data from the Excel calibration summary sheets into the boxes adjacent to the isotope line label.








Enter the ratio and uncertainty data if applicable.








If completely satisfied that the data has been correctly entered, write the new version (name extension) number and the changes made to the calibration in the comments box at the base of the screen, e.g. “103 – calibration for Th updated, 15 May 2010 (JP). Calibration OK”








Enter the new calibration name in the text box provided under ‘Save As’- follow the format of previous calibration files for that detector and geometry and increment the version, or extension number, by 1, eg next calibration after Cal9n.102 is Cal9n.103.








Double check that the correct detector, geometry and data have been selected or entered and then click on the ‘Confirm Changes’ button. This new calibration file is now the active one for that detector and geometry.








Click on ‘CLOSE’ to complete calibration entry.








This completes the validation process for that calibration sheet.





The latest calibrations must also be updated on the Sharepoint directory:


SSDX Equipment and Techniques > HPGe Gamma Spec > Calibrations > (year).








The printout is stored in the Gamma Detectors Calibration folder located in the Count Room, L29. 








If all calibration changes are checked as ‘OK’ on the printout, the person doing the checking validates the calibration changes by writing ‘Validated’ and signing and dating the printout.





Each changed calibration data value viewed on the computer is checked against the printed version. 


Usually this is done by two people, one calling the values from the computer and the other confirming the correct value and marking or ticking off the corresponding values on the printout.








Open Visual Gamma and click on the ‘View Detector Settings’ button. Select the ‘Detector’ and ‘Geometry’ from the drop-down menus.





Print out the Excel summary sheet containing the calibration data requiring validation. 














ii
iv

_1338100615.unknown

_1341817391.unknown

_1341817613.unknown

_1251797414.unknown

_1251879681.unknown

_1252138593.unknown

_1251797525.unknown

_1251792868.unknown

