# Vapour heat treatment sensor calibration record

Establishment name: ....................………………………………………………………………………………………………………………………

Establishment number: ....................……………………………………………………………………………………………………………………

NATA certified Mercury in Glass (MiG) reference thermometer used?

Yes ☐	Date of NATA certification (must be certified within last 3 years): ………………………………………………………..

No ☐ 	Reference thermometer not valid

Sensor probes to be inserted into vat of hot water, temperature 47 °C or more.  Authorised Officer to verify probe readings against MiG and ensure a variance of no greater than 0.3 °C (or 0.2 degrees for China). Each probe must be calibrated 3 times.

- First calibration all probes met requirements

Yes ☐ 			No ☐

- Second calibration all probes met requirements

Yes ☐ 			No ☐

- Third calibration all probes met requirements

Yes ☐ 			No ☐

Note: Attach the computer print-out of the probe calibration data and MiG thermometer NATA test certificate.

Probe/s that do not meet the requirements are listed below and cannot be used for VHT treatment:

…………………………………………………………………………………………………………………………………………………..………

Comments/corrective action: ………………………………………………………………………………………………….…………

……………………………………………………………………………………………………………………………………………….….……….

Authorised Officer name and number: ...........……………………............…....................................................................

Signature: .............……………………………………..……………………………………………	Date: ……………………………………..