# VIDAS Immuno Concentration Salmonella (ICS) - AOAC 2001.09

## SCOPE

This method is applicable to:

- Raw meats, meat products and carcass swabs

## PRINCIPLES

VIDAS ICS is a fully automated system. It replaces selective enrichment with an immunological capture and specific release procedure. This method should be read with reference to VIDAS SLM (AOAC 996.08). The detection of Salmonella in samples can be broken down as follows:

### Pre-enrichment in non-selective liquid medium (meat and meat products)

Sample is diluted in pre-warmed buffered peptone water and pre-enriched at 37 ± 1C for 18-24 h. Raw poultry samples are pre-enriched in BPW at 41 ± 1C for 18-24 h. For carcass sponges, buffered peptone water is added to the moistened sponge to bring the total volume to 60-100 mL and the sample incubated at 37 ± 1C for 18-24 h. In the case of sponges BPW need not be warmed to room temperature before being used to re-hydrate the sponge, for all subsequent additions BPW should be warmed to room temperature.

### Immuno-concentration

Salmonella is concentrated from pre-enrichment broth using the VIDAS ICS assay (following the manufacturer’s instructions).

### Post-enrichment

Within 30-minutes of the ICS assay the sample is enriched in ICS broth at 41 ± 1C for 5-6 h.

### Immuno-screening assay

Salmonella is assayed using VIDAS Salmonella (SLM) by testing a portion of boiled ICS broth.

### Cultural confirmation

All Salmonella positive samples should be confirmed using AS 5013.10. Confirmation must be carried out according to the manufacturer’s instructions using AS 5013.10 as appropriate, if confirmation is carried out off-site then the ‘pre-enrichment’ broth (i.e. BPW) must be used.

## CHECKLIST

| Pre-enrichment | Is the buffered peptone water (BPW) used for pre-enrichment? |  |
| --- | --- | --- |
|  | Is BPW warmed prior to use? |  |
|  | Is the correct amount of enrichment broth used for the weight of sample analysed? |  |
|  | Is a positive control run with each batch of samples analysed? |  |
|  | Are reference cultures inoculated into primary enrichment broth at a level of 10 to 100 cells? |  |
|  | Is pre-enrichment done at 37 ± 1C (41 ± 1C for raw poultry) for 18-24 h? |  |
| Post-Enrichment | Are the manufacturer’s instructions available? |  |
|  | Is post-enrichment initiated within 30-minutes of capture? |  |
|  | Is post-enrichment done at 41 ± 1C for 5-6 h? |  |
| Immuno-screening | Is the VIDAS SLM assay performed on boiled ICS broth? |  |
|  | Are the manufacturer’s instructions available and are they reproduced in the laboratory manual? |  |
|  | Are kites stored at 2-8C when not in use? |  |
| Cultural confirmation | Is Salmonella confirmed from pre-enrichment broth (if other, give details)? |  |
| (if applicable) | Are Salmonella confirmed using AS 5013.10 (with appropriate selective agar plates)? |  |
|  | If an external laboratory is used is it department approved? |  |
|  | BPW should be supplied to off-site laboratories for confirmation following AS 5013.10 |  |