# VIDAS® EASY Salmonella method - AFNOR BIO 12/16 – 09/05

## SCOPE

This method is applicable to:

- Raw meats, meat products and carcass swabs.

## PRINCIPLES

Detection of Salmonella is based on enzyme-linked fluorescent immunoassay performed in an automated VIDAS instrument. Following enrichment (see below) Salmonella is detected in boiled broth by specific monoclonal antibodies. A fluorescent marker is then added and Salmonella detected automatically by the VIDAS machine. The detection protocol can be broken down as follows:

### Pre-enrichment

Sample is diluted 1:10 in pre-warmed (room temperature or 37C for large volumes) buffered peptone water and pre-enriched at 37C for 16 - 22 h. For carcass sponges, buffered peptone water is added to the moistened sponge to bring the total volume to 60-100 ml and the sample incubated at 37C for 16 - 22 h. In the case of sponges BPW need not be warmed to room temperature before being used to re-hydrate the sponge, for all subsequent additions BPW should be warmed to room temperature.

### Enrichment

Culture from pre-enrichment broth is inoculated into SX2 broth and incubated at 41.5C for 22 - 26 h.

### Enzyme immunoassay

Salmonella is assayed by testing a portion of boiled post-enrichment broth. The assay is automated and carried out in the VIDAS instrument following the manufacturer’s instructions.

### Cultural confirmation

All Salmonella positive samples should be confirmed using the stored non heated SX2 broth following AS 5013.10. Confirmation carried out at an ‘off-site’ laboratory must be from retained SX2 enrichment.

## CHECKLIST

| Pre-enrichment | Is the buffered peptone water used for pre-enrichment? |  |
| --- | --- | --- |
|  | Is BPW warmed prior to use (room temperature or 37C)? |  |
|  | Is the correct amount of enrichment broth used for the weight of sample analysed? |  |
|  | Is a positive control run with each batch of samples analysed? |  |
|  | Are reference cultures inoculated into primary enrichment broth at a level of 10 to 100 cells? |  |
|  | Is pre-enrichment done at 37C for 16 - 22h? |  |
| Enrichment | Is enrichment carried out in SX2 broth? |  |
|  | Is SX2 broth incubated at 41.5C for 22- 26h? |  |
| Enzyme immunoassay | Are the manufacturer’s instructions available? |  |
|  | Are kites stored at 2-8 C when not in use? |  |
| Cultural confirmation | Is Salmonella confirmed from SX2 Broth? |  |
| (if applicable) | If an external laboratory is used is it department approved? |  |
|  | SX2 enrichment should be supplied to off-site laboratories for confirmation following AS 5013.10 |  |
|  | Are Salmonella confirmed using AS 5013.10 (with appropriate selective agar plates)? |  |